## Alarm sheet
File name | HCMB__A001-02000018_cc.xls
Component ID | HCMB__A001 |  |  |  | Serial Number | 2000018
Date of test Ap 1 | 17/10/2002 |  |  |  | Date of test Ap 2 | 17/10/202
| Aperture 1 |  |  |  |  | Aperture 2
Magnetic length | status ok |  |  |  | Magnetic length | status ok
| Average straight | Variation straight | Heads CS | Heads NCS |  | Average straight | Variation straight | Heads CS | Heads NCS
| positions 2 to 19 | positions 2 to 19 | position 1 | position 20 |  | positions 2 to 19 | positions 2 to 19 | position 1 | position 20
Main field | status ok | status ok | status ok | status ok | Main field | status ok | status ok | status ok | status ok
Angle |  | status ok | status ok | status ok | Angle |  | status ok | status ok | status ok
b2 | status ok | status ok | status ok | status ok | b2 | status ok | status ok | status ok | status ok
b3 | status ok | status ok | status ok | yellow alarm | b3 | status ok | status ok | status ok | yellow alarm
b4 | status ok | status ok | status ok | status ok | b4 | status ok | status ok | status ok | status ok
b5 | status ok | status ok | status ok | status ok | b5 | status ok | status ok | status ok | status ok
b6 | status ok | status ok | status ok | status ok | b6 | status ok | status ok | status ok | status ok
b7 | status ok | status ok | status ok | status ok | b7 | status ok | status ok | status ok | status ok
b8 | status ok | yellow alarm | status ok | status ok | b8 | status ok | yellow alarm | status ok | status ok
b9 | status ok | status ok | status ok | status ok | b9 | status ok | status ok | status ok | status ok
b10 | status ok | yellow alarm | status ok | status ok | b10 | status ok | yellow alarm | status ok | status ok
b11 | status ok | status ok | status ok | status ok | b11 | status ok | status ok | status ok | status ok
b12 | status ok | yellow alarm | status ok | status ok | b12 | status ok | yellow alarm | status ok | status ok
b13 | status ok | status ok | status ok | status ok | b13 | status ok | status ok | status ok | status ok
b14 | status ok | status ok | status ok | status ok | b14 | status ok | status ok | status ok | status ok
b15 | status ok | yellow alarm | status ok | status ok | b15 | status ok | yellow alarm | status ok | status ok
a2 | status ok | status ok | status ok | status ok | a2 | status ok | status ok | status ok | status ok
a3 | status ok | status ok | status ok | status ok | a3 | status ok | status ok | status ok | status ok
a4 | status ok | status ok | status ok | status ok | a4 | status ok | status ok | status ok | status ok
a5 | status ok | status ok | status ok | status ok | a5 | status ok | status ok | status ok | status ok
a6 | status ok | status ok | status ok | status ok | a6 | status ok | status ok | status ok | status ok
a7 | status ok | status ok | status ok | status ok | a7 | status ok | status ok | status ok | status ok
a8 | status ok | status ok | status ok | status ok | a8 | status ok | yellow alarm | status ok | status ok
a9 | status ok | status ok | status ok | status ok | a9 | status ok | status ok | status ok | status ok
a10 | status ok | status ok | status ok | status ok | a10 | status ok | yellow alarm | status ok | status ok
a11 | status ok | status ok | status ok | status ok | a11 | status ok | status ok | status ok | yellow alarm
a12 | status ok | status ok | status ok | status ok | a12 | status ok | status ok | status ok | status ok
a13 | status ok | status ok | status ok | status ok | a13 | status ok | status ok | status ok | status ok
a14 | status ok | status ok | status ok | status ok | a14 | status ok | status ok | status ok | status ok
a15 | status ok | status ok | status ok | status ok | a15 | status ok | status ok | status ok | status ok
Coil Positioning | status ok |  |  |  | Coil Positioning | status ok
Field Colinearity | status ok

## Assembly Data
Prog. # cc | Magnet name |  |  | Coil cross-section |  |  | End Spacers
| HCMB__A001 | 2000018 |  | 2 |  |  | 3
COMMENTS:
| HOLDING POINT |  | Data at CERN |  |  | Analysis by |  |  | Answer to MMS-LD |  |  |  | Answer to manufacturer
| Shims-summary | [6 empty] | Shims
| Inner average |  |  | Outer average |  |  |  | Aperture 1 |  |  |  |  | Aperture 2
| Used | Nominal | Differ. | Used | Nominal | Differ. |  | Left |  | Right |  |  | Left |  | Right
| 0.1 | 0.2 | -0.1 | 0.8 | 0.8 | 0 |  | Outer | Inner | Inner | Outer |  | Outer | Inner | Inner | Outer
[8 empty] | 0.8 | 0.1 | 0.1 | 0.8 | Up | 0.8 | 0.1 | 0.1 | 0.8
| Pre-stress (estimate, Mpa) | [6 empty] | 0.8 | 0.1 | 0.1 | 0.8 | Down | 0.8 | 0.1 | 0.1 | 0.8
| Inner average |  |  | Outer average |  |  |  | Seen from conection side
| Used | Nominal | Differ. | Used | Nominal | Differ.
| 75 | 75 | 0 | 75 | 75 | 0
| Peak pre-stress during collaring |  |  |  |  | 110
| Collared coil template - version 11.07.2002 | [8 empty] | Version : |  | cern
1 | Remplir les champs jaunes dans "Assembly Data" et dans "Original Data" | [12 empty] | French
2 | CTRL-g pour charger les fichers des mesures .txt
3 | Sauver le fichier avec un nom different
4 | ALT-F8 et appuyer sur "collared coil analysis" pour analyser les mesures - resultats dans Alarm sheet
1 | Fuellen Sie die gelben Felder in "Assembly Data" und in "Original Data" | [12 empty] | German
2 | CTRL-g um die .txt Messdatensaetze zu laden
3 | Sichern Sie diesen Datensatz mit einem anderen Namen
4 | ALT-F8 und auf "collared coil analysis" druecken, um die Daten zu analysieren => Ergebnis in "Alarm sheet"
1 | Riempire I campi gialli in "Assembly Data" e in "Original Data" | [12 empty] | Italian
2 | CTRL-g per caricare I file .txt delle misure
3 | Salvare il file con un nome diverso
4 | ALT-F8 e premere su "collared coil analysis" per analiszzare le misure - risultati in "Alarm sheet"
1 | Fill in the yellow fields in "Assembly Data" and "Original Data" | [12 empty] | English
2 | CTRL-g to load the .txt files of the measurements
3 | Save the file with a different name
4 | ALT-F8 and press on "collared coil analysis" to analyse the measures - results in Alarm sheet
| Last modif on | by | Modifications:
| 2002-07-10 | E. Todesco | New template with limits for X-section 2

## Original data
File name |  | HCMB__A001-02000018_cc.xls
Component ID |  | HCMB__A001 |  |  | Serial Number |  |  | 2000018 | [11 empty] | Signs
Ref. Test Proced. |  | CERN IT 2708/LHC/LHC Rev 1.1 Annex b.18 | [16 empty] | b even | 1
Firm Name |  | Ansaldo | [9 empty] | File name for aperture 1 |  |  | Ansaldo18cct1.txt |  |  |  | a odd | 1
Test Type |  | Magnetic Measures |  |  | ITP number |  |  | 21 |  |  |  | File name for aperture 2 |  |  | Ansaldo18cct2.txt |  |  |  | a even | 1
Test Operator |  | A. Musso |  |  | Test Controller |  |  | E. Todesco
Aperture 1 | [11 empty] | Aperture 2
Date of test |  | 17/10/2002 |  |  | Time of test |  |  | 10:00 - 13:00 |  |  |  | Date of test |  | 17/10/202 |  |  | Time of test |  |  | 14:00 - 17:00
Mole name |  | Dipole 6 |  |  | Calibration date |  |  | 22/7/2002 |  |  |  | Mole name |  | Dipole 6 |  |  | Calibration date |  |  | 22/07/2002
Coil length (m) |  | 0.7499 |  |  | Coil temperature |  |  | 23 |  |  |  | Coil length (m) |  | 0.7499 |  |  | Coil temperature |  |  | 23.5
Excitation current (A) |  | 8.5 |  |  | Analysis tools |  |  | dri,rot,nor,cel,fdw |  |  |  | Excitation current (A) |  | 8.5 |  |  | Analysis tools |  |  | dri,rot,nor,cel,fdw
N. points |  | 15 |  |  | N. measure |  |  | 10 (5+,5-) |  |  |  | N. points |  | 15 |  |  | N. measure |  |  | 10 (5+,5-)
Rotation speed |  | 0.6 |  |  | I. Gain |  |  | 50/500 |  |  |  | Rotation speed |  | 0.6 |  |  | I. Gain |  |  | 50/500
Dx average |  | 0.00023315 |  |  | Dy average |  |  | 0.00029042 |  |  |  | Dx average |  | 0.00020586 |  |  | Dy average |  |  | 0.00040987
Magnetic length |  | 14.4479 |  |  | Tranfs. Function (mT/KA) |  |  | 595.0719 |  |  |  | Magnetic length |  | 14.4484 |  |  | Tranfs. Function (mT/KA) |  |  | 595.0196
Phase |  | -3.9103 |  |  | Integral field (T.m) |  |  | 0.073079 |  |  |  | Phase |  | -3.5802 |  |  | Integral field (T.m) |  |  | 0.073074
File | Aperture 1 - Collared coils | [19 empty] | Integrals | File | Aperture 2 - Collared coils | [19 empty] | Integrals
C1 (T) | 0.003213 | 0.005061 | 0.005058 | 0.005058 | 0.005057 | 0.005057 | 0.005057 | 0.005057 | 0.005058 | 0.005058 | 0.005058 | 0.005058 | 0.005058 | 0.005058 | 0.005059 | 0.005058 | 0.005057 | 0.005058 | 0.005061 | 0.003192 | 0.073079 | C1 (T) | 0.003218 | 0.005061 | 0.005058 | 0.005057 | 0.005057 | 0.005057 | 0.005056 | 0.005056 | 0.005058 | 0.005057 | 0.005057 | 0.005057 | 0.005057 | 0.005058 | 0.005058 | 0.005057 | 0.005058 | 0.005058 | 0.005061 | 0.00319 | 0.073074
Angle (mrad) | 1.8824 | -0.82482 | -0.6497 | -0.11477 | -0.25935 | 0.014966 | -0.14398 | -0.12653 | 0.090031 | 0.013652 | -0.51237 | -0.19961 | 0.2163 | -0.67277 | -0.42955 | 0.49493 | 0.50844 | 0.54944 | 0.59499 | 0.43857 | -3.9103 | Angle (mrad) | 3.5224 | 0.36162 | 0.63676 | -0.24642 | -1.0256 | -0.22761 | -0.18658 | 0.59754 | 0.53253 | -0.74684 | -0.61359 | -0.16414 | -0.40797 | -0.53623 | -0.38007 | -0.28128 | -0.10488 | -0.006099 | 0.35609 | 0.31202 | -3.5802
Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20 |  | Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20
b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000
b2 | -2.0638 | -0.051021 | -0.4426 | 0.13584 | 0.70615 | 0.349 | 0.43686 | 1.0856 | 0.27648 | -0.21222 | 0.67978 | 0.23684 | 0.56453 | 0.92054 | 0.86058 | 0.31889 | 0.45938 | 0.88595 | 0.9089 | 0.14535 | 0.35813 | b2 | -4.7971 | 0.20419 | -0.14349 | -1.5695 | -0.49505 | -0.73748 | -0.16189 | -0.9544 | -0.79242 | -1.7941 | -1.4567 | -1.5843 | -1.0641 | -0.20739 | 0.051171 | -0.022432 | -1.5222 | -1.098 | 0.13367 | -0.72161 | -0.86781
b3 | 26.0355 | -2.1618 | -2.8936 | -2.7467 | -2.9981 | -3.0902 | -3.2382 | -3.1072 | -2.7766 | -3.188 | -3.0673 | -2.8117 | -2.6828 | -3.0028 | -2.8507 | -2.772 | -2.7103 | -2.8927 | -2.5156 | -5.2297 | -1.9861 | b3 | 28.936 | -2.0992 | -2.2896 | -2.0595 | -2.5694 | -2.7214 | -2.8192 | -2.3549 | -2.0073 | -2.9357 | -2.4187 | -2.4501 | -2.0415 | -1.9595 | -2.0864 | -2.3981 | -1.69 | -1.8963 | -1.5172 | -4.6058 | -1.2876
b4 | -1.3914 | 0.053072 | -0.029873 | -0.023156 | 0.012658 | -0.076898 | -0.049485 | 0.024348 | 0.086402 | -0.0026736 | -0.03115 | -0.038478 | -0.038093 | 0.042586 | -0.023974 | -0.057243 | 0.002847 | -0.030678 | 0.059237 | -0.29276 | -0.061722 | b4 | -0.74965 | -0.076509 | -0.010478 | -0.10808 | -0.057437 | -0.20661 | -0.23301 | -0.074624 | -0.064175 | -0.044832 | -0.21625 | -0.14461 | -0.070989 | -0.19913 | 0.010734 | -0.039225 | 0.076583 | -0.0092453 | 0.013693 | 0.012733 | -0.099805
b5 | -5.7754 | 0.14498 | 0.046533 | 0.060776 | 0.061861 | 0.34342 | 0.2787 | 0.25658 | 0.2691 | 0.32329 | 0.21441 | 0.1663 | 0.25166 | 0.22998 | 0.21043 | 0.15811 | 0.26815 | 0.15484 | 0.12417 | -4.1581 | -0.1417 | b5 | -5.7775 | 0.18639 | 0.20349 | 0.26972 | 0.35528 | 0.53981 | 0.19299 | 0.21942 | 0.27357 | 0.36161 | 0.25295 | 0.20006 | 0.2707 | 0.2466 | 0.32042 | 0.12679 | 0.25216 | 0.12273 | -0.071807 | -3.4449 | -0.079228
b6 | 0.50395 | 0.11768 | 0.056449 | -0.085771 | 0.090026 | 0.075011 | 0.082149 | -0.0097875 | 0.11242 | 0.15104 | -0.038485 | 0.10506 | -0.068075 | 0.01361 | -0.016817 | 0.094263 | -0.057394 | -0.018418 | 0.10032 | -0.0050193 | 0.052961 | b6 | -0.10598 | -0.13323 | -0.066877 | -0.067736 | -0.11701 | -0.11582 | -0.046592 | -0.2221 | -0.085561 | -0.025786 | -0.06414 | -0.01053 | -0.12075 | -0.075465 | -0.054078 | -0.0061507 | -0.10958 | -0.096624 | 0.055391 | -0.010278 | -0.074557
b7 | 1.6647 | 0.72354 | 0.82844 | 0.85758 | 0.82814 | 0.80357 | 0.87359 | 0.81673 | 0.83872 | 0.79825 | 0.82612 | 0.82137 | 0.86092 | 0.84236 | 0.77371 | 0.83755 | 0.8747 | 0.82469 | 0.82898 | 0.2716 | 0.83502 | b7 | 1.7078 | 0.7626 | 0.89842 | 0.87911 | 0.91606 | 0.85686 | 0.90335 | 0.87034 | 0.88759 | 0.8446 | 0.88488 | 0.89482 | 0.90973 | 0.9196 | 0.87183 | 0.8712 | 0.84397 | 0.82782 | 0.84553 | 0.17828 | 0.87648
b8 | 0.096574 | 0.064552 | 0.0291 | -0.14824 | 0.023236 | -0.037041 | 0.0062094 | -0.021231 | 0.042073 | 0.051612 | -0.085887 | 0.064947 | -0.071656 | 0.036849 | 0.024181 | 0.07202 | -0.032797 | 0.042885 | 0.067573 | -0.010899 | 0.0094948 | b8 | -0.071928 | 1.8514e-05 | -0.04836 | 0.016633 | -0.02818 | -0.0049359 | -0.028799 | -0.1479 | -0.039535 | 0.070896 | 0.0098976 | 0.034643 | -0.037551 | -0.009066 | 0.01907 | 0.019183 | 0.004589 | 0.0050925 | 0.042389 | -0.097212 | -0.011883
b9 | 0.44198 | 0.60334 | 0.57554 | 0.60148 | 0.59081 | 0.58578 | 0.58998 | 0.59568 | 0.59585 | 0.59012 | 0.59696 | 0.59051 | 0.58937 | 0.58617 | 0.58901 | 0.58765 | 0.59708 | 0.57843 | 0.59939 | 0.45432 | 0.58188 | b9 | 0.39137 | 0.60872 | 0.60379 | 0.59753 | 0.59915 | 0.58917 | 0.59139 | 0.59596 | 0.58717 | 0.6028 | 0.60419 | 0.60162 | 0.60765 | 0.58776 | 0.60951 | 0.59129 | 0.60857 | 0.60173 | 0.60415 | 0.4823 | 0.58885
b10 | 0.20918 | 0.068826 | 0.071918 | -0.2289 | 0.062297 | -0.039221 | 0.024303 | -0.074146 | 0.063381 | 0.12219 | -0.18217 | 0.080004 | -0.14693 | 0.038273 | -0.0041698 | 0.004281 | -0.11152 | 0.0091031 | 0.075186 | 0.063669 | 0.00030556 | b10 | 0.035514 | -0.027497 | -0.10609 | 0.023904 | -0.05224 | 0.029771 | -0.011741 | -0.20766 | -0.048031 | 0.10005 | 0.038202 | 0.059669 | -0.066174 | 0.015965 | 0.050363 | 0.069233 | 0.010831 | 0.046326 | 0.10153 | -0.054805 | 0.00075281
b11 | 0.60687 | 0.77097 | 0.76696 | 0.7653 | 0.76476 | 0.76557 | 0.76162 | 0.76515 | 0.76736 | 0.77387 | 0.77388 | 0.77111 | 0.76972 | 0.76857 | 0.77279 | 0.76546 | 0.76571 | 0.76303 | 0.76319 | 0.6318 | 0.75776 | b11 | 0.58659 | 0.76351 | 0.7648 | 0.76143 | 0.76575 | 0.76533 | 0.7586 | 0.76016 | 0.76202 | 0.76619 | 0.76064 | 0.76021 | 0.76196 | 0.7606 | 0.7685 | 0.76256 | 0.77065 | 0.76466 | 0.7655 | 0.65528 | 0.75412
b12 | 0.036897 | 0.0080949 | 0.008208 | -0.032985 | 0.0075716 | -0.0075935 | 0.0019986 | -0.0061723 | 0.0057549 | 0.011095 | -0.023297 | 0.011522 | -0.019408 | 0.0073324 | 0.0031759 | 0.0023788 | -0.013636 | 0.0031129 | 0.011591 | 0.0049555 | 0.00027651 | b12 | -0.0048772 | 0.0015772 | -0.011663 | 0.0043422 | -0.0053051 | 0.007839 | -0.00012234 | -0.028769 | -0.0069005 | 0.013545 | 0.0038607 | 0.0066638 | -0.0076473 | 0.0040446 | 0.0096194 | 0.011859 | 0.00020148 | 0.0041449 | 0.014043 | -0.014588 | 0.00046927
b13 | 0.083063 | 0.082549 | 0.085627 | 0.089809 | 0.086091 | 0.083808 | 0.084066 | 0.081457 | 0.083408 | 0.083469 | 0.087019 | 0.084998 | 0.086742 | 0.084035 | 0.084409 | 0.084345 | 0.086214 | 0.082064 | 0.083091 | 0.052939 | 0.083533 | b13 | 0.082833 | 0.084396 | 0.087665 | 0.084472 | 0.08491 | 0.081831 | 0.087518 | 0.086438 | 0.08648 | 0.08381 | 0.086276 | 0.087409 | 0.086614 | 0.086693 | 0.086861 | 0.087477 | 0.084453 | 0.085033 | 0.089793 | 0.047484 | 0.084641
b14 | 0.0020249 | -0.0014473 | -0.0019614 | -0.012649 | -0.0019154 | -0.0074148 | -0.0036171 | -0.0070546 | -0.0015036 | 0.0009605 | -0.012065 | -0.00061659 | -0.010079 | -0.0018511 | -0.0045711 | -0.0042419 | -0.009659 | -0.0045697 | -0.00033904 | -0.0037957 | -0.0044482 | b14 | 0.0024164 | -0.0055835 | -0.0078279 | -0.0031159 | -0.0073449 | -0.0016444 | -0.0033635 | -0.010685 | -0.0046585 | 0.0013511 | -0.0015408 | -0.00060898 | -0.0058143 | -0.0016622 | -0.0015738 | 0.00035681 | -0.0045269 | -0.0018462 | 0.0017865 | -0.0057776 | -0.0031352
b15 | 0.002472 | 0.030844 | 0.027607 | 0.01162 | 0.027521 | 0.026246 | 0.027923 | 0.023759 | 0.028244 | 0.032047 | 0.016833 | 0.029087 | 0.019373 | 0.026507 | 0.025469 | 0.028071 | 0.024626 | 0.031381 | 0.028365 | 0.014158 | 0.024708 | b15 | -0.00086662 | 0.030541 | 0.02344 | 0.032495 | 0.028181 | 0.033943 | 0.027796 | 0.013701 | 0.023664 | 0.029349 | 0.027297 | 0.028012 | 0.024263 | 0.028546 | 0.028999 | 0.028022 | 0.030721 | 0.030855 | 0.027076 | 0.014979 | 0.026252
b16 | [21 empty] | b16
b17 | [21 empty] | b17
a1 | 18.8241 | -8.2482 | -6.497 | -1.1477 | -2.5935 | 0.14966 | -1.4398 | -1.2653 | 0.90031 | 0.13652 | -5.1237 | -1.9961 | 2.163 | -6.7277 | -4.2955 | 4.9493 | 5.0844 | 5.4944 | 5.9499 | 4.3857 | 0 | a1 | 35.2246 | 3.6162 | 6.3676 | -2.4642 | -10.2557 | -2.276 | -1.8658 | 5.9754 | 5.3253 | -7.4684 | -6.1359 | -1.6414 | -4.0797 | -5.3623 | -3.8007 | -2.8128 | -1.0488 | -0.060989 | 3.561 | 3.1202 | -1.0386e-13
a2 | 9.5164 | -0.21749 | 0.26167 | -0.28232 | 0.11671 | 0.19313 | 0.84531 | 1.3531 | 1.0528 | 0.61795 | -0.094188 | 0.15876 | 0.61673 | 0.34182 | 0.14146 | -0.42609 | 0.48546 | 0.016364 | 0.22346 | -9.9728 | 0.26762 | a2 | 9.3715 | 1.3547 | 1.2407 | 1.4723 | 1.3681 | 1.0542 | 1.4179 | 2.0407 | 1.3041 | 1.1769 | 1.5194 | 1.3247 | 1.5278 | 1.3784 | 1.3027 | 1.5316 | 1.8122 | 1.541 | 2.1006 | -10.3166 | 1.3455
a3 | -3.0212 | -0.64295 | -0.85956 | -0.77732 | -0.84449 | -1.1531 | -1.3503 | -0.39723 | -0.50729 | -0.65813 | -0.23471 | -0.39993 | -0.43822 | -0.61129 | -0.64941 | -0.57639 | -1.0639 | -0.68675 | -0.50774 | 0.14103 | -0.73644 | a3 | -0.69418 | -0.17595 | -0.4763 | -0.52938 | -0.16154 | -0.47723 | -0.49568 | -0.35228 | -0.23583 | -0.52241 | 0.092318 | -0.042028 | 0.087491 | -0.044304 | -0.87355 | -1.0217 | -0.63041 | -0.61542 | -1.0455 | -0.88559 | -0.44223
a4 | 2.6065 | 0.20607 | 0.012266 | 0.21055 | 0.079676 | 0.38885 | 0.26607 | 0.3062 | 0.22162 | -0.029672 | -0.01609 | 0.1165 | 0.11908 | 0.030194 | 0.07755 | 0.12751 | 0.33749 | 0.055272 | 0.051359 | -0.15436 | 0.21378 | a4 | 1.5779 | 0.16592 | 0.085279 | 0.25612 | 0.63282 | 0.49678 | 0.046485 | 0.12301 | 0.22404 | 0.37541 | 0.42668 | 0.49564 | 0.41989 | 0.60952 | 0.34375 | 0.28211 | 0.26851 | 0.25218 | 0.12654 | -0.7786 | 0.31885
a5 | 1.3367 | -0.1318 | -0.18704 | -0.12701 | -0.15994 | -0.14913 | -0.25722 | -0.084677 | -0.028439 | -0.093211 | -0.13761 | -0.08878 | -0.12281 | -0.08068 | -0.13896 | -0.20392 | -0.17281 | -0.059998 | -0.11988 | -0.58608 | -0.096777 | a5 | 1.6208 | -0.092331 | -0.13483 | -0.17549 | -0.12986 | -0.032703 | -0.16782 | -0.1241 | -0.037571 | -0.17 | -0.087298 | -0.016137 | -0.024566 | 0.048223 | -0.19417 | -0.18492 | -0.14461 | -0.19763 | -0.29091 | -0.56584 | -0.076943
a6 | 0.033892 | 0.042485 | 0.089843 | 0.16762 | 0.093949 | 0.097656 | 0.10038 | 0.21261 | 0.13581 | 0.091547 | 0.15311 | 0.072923 | 0.17653 | 0.13125 | 0.083411 | 0.034913 | 0.046267 | 0.021004 | 0.049667 | -0.13684 | 0.090109 | a6 | -0.18941 | 0.024543 | -0.071728 | -0.087507 | 0.03095 | -0.025507 | -0.097404 | 0.0091237 | 0.011617 | -0.021017 | 0.0082439 | 0.0070975 | 0.07439 | 0.089158 | -0.046071 | -0.028279 | -0.088285 | -0.013156 | 0.041512 | 0.10453 | -0.012291
a7 | 1.7701 | 0.1263 | 0.071607 | 0.016504 | 0.024684 | 0.051631 | 0.044706 | 0.052753 | 0.10479 | 0.081492 | 0.017245 | 0.058248 | 0.047839 | 0.065118 | 0.064476 | 0.070608 | 0.045424 | 0.04808 | 0.052644 | 0.070396 | 0.11487 | a7 | 1.8741 | 0.073127 | 0.048956 | 0.049624 | 0.075129 | 0.16779 | 0.094429 | 0.063959 | 0.147 | 0.12426 | 0.095179 | 0.10988 | 0.13575 | 0.10763 | 0.028654 | 0.025213 | 0.060996 | 0.0644 | 0.03727 | 0.17106 | 0.14582
a8 | -0.046997 | -0.018795 | 0.021269 | 0.084949 | -0.014737 | -0.011417 | -0.027318 | -0.042814 | -0.032886 | -0.070276 | 0.038937 | -0.037639 | 0.070183 | 0.018161 | 0.016738 | -0.061717 | -0.042416 | -0.089627 | -0.062737 | -0.074966 | -0.017612 | a8 | -0.089879 | -0.025956 | 0.018581 | -0.032205 | 0.026751 | 0.00076902 | -0.017626 | 0.1217 | 0.067961 | 0.02269 | 0.040408 | -0.0028055 | 0.054542 | 0.025565 | 0.016873 | 0.0071511 | -0.0022602 | 0.022043 | 0.031192 | -0.049248 | 0.014901
a9 | -0.31141 | 0.05622 | 0.049165 | 0.0011294 | 0.040037 | 0.047633 | 0.04023 | 0.062387 | 0.046063 | 0.030133 | 0.05098 | 0.061108 | 0.046295 | 0.047965 | 0.038758 | 0.079508 | 0.081143 | 0.088852 | 0.05942 | 0.018806 | 0.038458 | a9 | -0.22032 | 0.082933 | 0.026725 | 0.041548 | 0.017758 | 0.056892 | 0.046715 | 0.024071 | 0.055418 | 0.015928 | 0.069863 | 0.042548 | 0.05396 | 0.033435 | 0.035114 | 0.056488 | 0.071083 | 0.067204 | 0.05109 | -0.040857 | 0.035442
a10 | -0.069892 | -0.042217 | 0.052305 | 0.11582 | -0.010043 | -0.0085898 | -0.023735 | -0.00086131 | -0.015264 | -0.064933 | 0.10853 | -0.046303 | 0.12437 | 0.053991 | 0.082337 | -0.021684 | -0.043931 | -0.10949 | -0.042115 | -0.10569 | -0.00015175 | a10 | -0.092078 | -0.020164 | 0.034508 | -0.079625 | -0.018544 | -0.056363 | -0.005284 | 0.17526 | 0.093067 | 0.00023885 | 0.014829 | -0.0135 | 0.03735 | -0.010655 | -0.0027446 | 0.014994 | -0.039842 | -0.041352 | 0.036689 | -0.08841 | 0.00023329
a11 | 0.25846 | 0.066763 | 0.060337 | 0.057093 | 0.060555 | 0.068528 | 0.065639 | 0.05988 | 0.073527 | 0.065503 | 0.050247 | 0.061879 | 0.060998 | 0.059727 | 0.06125 | 0.06947 | 0.070457 | 0.071832 | 0.06691 | 0.06824 | 0.070469 | a11 | 0.26454 | 0.067539 | 0.067071 | 0.058963 | 0.053588 | 0.077101 | 0.06355 | 0.063191 | 0.081118 | 0.063377 | 0.060181 | 0.068477 | 0.06519 | 0.066187 | 0.059033 | 0.065462 | 0.063169 | 0.06044 | 0.064455 | 0.092424 | 0.07239
a12 | -0.023779 | -0.0044778 | 0.007767 | 0.017827 | 0.001096 | 0.0015517 | 0.00052619 | 0.0071679 | 0.0023061 | -0.0053362 | 0.016527 | -0.0048377 | 0.019278 | 0.0093526 | 0.010716 | -0.0021378 | -0.0030559 | -0.010951 | -0.003011 | -0.011596 | 0.0019662 | a12 | -0.017591 | -0.0034266 | -0.0015339 | -0.014358 | -0.0028037 | -0.0077385 | -0.0057616 | 0.017256 | 0.010541 | -0.0022136 | 0.00088651 | -0.00053896 | 0.0074294 | -0.00023425 | -0.003 | -0.001368 | -0.0077697 | -0.0076488 | 0.0036634 | -0.0038417 | -0.001673
a13 | -0.018959 | 0.0079825 | 0.0063781 | 0.0050896 | 0.0049767 | 0.0065018 | 0.0037638 | 0.0082212 | 0.0070324 | 0.0061333 | 0.0080651 | 0.0085992 | 0.0084025 | 0.0057603 | 0.0052233 | 0.0096693 | 0.010093 | 0.010322 | 0.0081698 | 0.0037508 | 0.006264 | a13 | -0.010616 | 0.012525 | 0.0080936 | 0.007268 | 0.004073 | 0.0078196 | 0.0062694 | 0.0077304 | 0.0094499 | 0.0043428 | 0.0086339 | 0.0087027 | 0.010323 | 0.0083164 | 0.0055749 | 0.0074085 | 0.0084807 | 0.0075356 | 0.0048276 | -0.0012332 | 0.0067394
a14 | 0.011186 | -0.0097026 | -0.0058373 | -0.0019979 | -0.0076412 | -0.0061844 | -0.0071838 | -0.0075132 | -0.0078995 | -0.012139 | -0.0028801 | -0.010197 | -0.0015907 | -0.0053737 | -0.0050282 | -0.010211 | -0.0070264 | -0.013958 | -0.010224 | -0.004007 | -0.0066443 | a14 | 0.0068378 | -0.011256 | -0.0071458 | -0.012836 | -0.0080703 | -0.010698 | -0.0094664 | -0.0019365 | -0.0043676 | -0.0076951 | -0.0065383 | -0.0085739 | -0.0061221 | -0.0084756 | -0.0084241 | -0.0084926 | -0.010596 | -0.010356 | -0.0072692 | -0.0039953 | -0.0076032
a15 | 0.0096899 | 0.0011188 | 0.0017969 | -5.4093e-05 | -9.2627e-05 | -0.0042947 | -0.0030199 | -0.0028188 | 0.0014178 | 0.0029712 | 0.00060038 | 0.0014137 | 0.0016765 | 0.0011644 | 0.0008212 | 0.00010219 | -0.0075168 | -0.0027869 | 0.0017949 | -0.0050318 | -0.00014162 | a15 | -0.0033195 | -0.0025908 | -0.0043197 | -0.0018335 | -0.005266 | -0.00075535 | -0.0016213 | 0.0028192 | 0.000607 | 0.0044577 | 0.0008156 | 0.002037 | -0.0016777 | -3.7486e-05 | -0.00042475 | 0.00030498 | -0.0013647 | 1.9787e-05 | 0.0013497 | -0.015458 | -0.0010039
a16 | [21 empty] | a16
a17 | [21 empty] | a17
Dx (m) | -0.00042543 | -0.00014263 | -0.00016749 | 0.00048656 | -0.00013586 | 8.8094e-05 | -4.9315e-05 | 0.00016388 | -0.00013592 | -0.00025454 | 0.00038309 | -0.00016711 | 0.00030085 | -9.3371e-05 | -5.1507e-06 | -5.095e-06 | 0.00025442 | 2.6588e-06 | -0.00015804 | -0.00013898 | 0.00023315 | Dx (m) | 1.4476e-05 | 6.4688e-05 | 0.00022734 | -3.9367e-05 | 0.00011828 | -5.2979e-05 | 2.7114e-05 | 0.00042887 | 8.4098e-05 | -0.00022051 | -8.7456e-05 | -0.00012966 | 0.00013949 | -3.3357e-05 | -0.00011029 | -0.00015606 | -1.6577e-05 | -9.5133e-05 | -0.0002307 | 0.00017113 | 0.00020586
Dy (m) | 0.00037697 | 0.00010544 | -0.00010276 | -0.00029357 | 3.3082e-05 | 1.1189e-05 | 5.7229e-05 | -1.0912e-05 | 4.684e-05 | 0.00016419 | -0.00026327 | 0.00011549 | -0.00029854 | -0.00011217 | -0.00018072 | 4.862e-05 | 7.4125e-05 | 0.00024368 | 0.00010767 | 0.00029938 | 0.00029042 | Dy (m) | 0.00026032 | 3.9175e-05 | -9.6641e-05 | 0.00018082 | 3.2892e-05 | 0.00013053 | 9.5699e-06 | -0.0004276 | -0.00021658 | 1.771e-05 | -2.6222e-05 | 4.1868e-05 | -9.5266e-05 | 2.6718e-05 | 1.4544e-05 | -2.0028e-05 | 8.9248e-05 | 9.9453e-05 | -6.2053e-05 | 0.00020522 | 0.00040987
Magnetic length (m) |  | 14.4479 | [19 empty] | Magnetic length (m) |  | 14.4484
Transfer funct. (mT/KA) |  | 595.0719 | [19 empty] | Transfer funct. (mT/KA) |  | 595.0196
Coil positioning (mm) |  | 0.013299 | [19 empty] | Coil positioning (mm) |  | 0.016155
Current (A) |  | 8.5 | [19 empty] | Current (A) |  | 8.5

## Summary Data
File | HCMB__A001-2000018 |  |  | Aperture 1 - Collared coils | [16 empty] | Integrals | File | HCMB__A001-2000018 |  |  | Aperture 2 - Collared coils | [16 empty] | Integrals
C1/i (mT/KA) | 378 | 595.4118 | 595.0588 | 595.0588 | 594.9412 | 594.9412 | 594.9412 | 594.9412 | 595.0588 | 595.0588 | 595.0588 | 595.0588 | 595.0588 | 595.0588 | 595.1765 | 595.0588 | 594.9412 | 595.0588 | 595.4118 | 375.5294 | 0.073079 | C1/i (mT/KA) | 378.5882 | 595.4118 | 595.0588 | 594.9412 | 594.9412 | 594.9412 | 594.8235 | 594.8235 | 595.0588 | 594.9412 | 594.9412 | 594.9412 | 594.9412 | 595.0588 | 595.0588 | 594.9412 | 595.0588 | 595.0588 | 595.4118 | 375.2941 | 0.073074
Angle (mrad) | 1.8824 | -0.82482 | -0.6497 | -0.11477 | -0.25935 | 0.014966 | -0.14398 | -0.12653 | 0.090031 | 0.013652 | -0.51237 | -0.19961 | 0.2163 | -0.67277 | -0.42955 | 0.49493 | 0.50844 | 0.54944 | 0.59499 | 0.43857 | -3.9103 | Angle (mrad) | 3.5224 | 0.36162 | 0.63676 | -0.24642 | -1.0256 | -0.22761 | -0.18658 | 0.59754 | 0.53253 | -0.74684 | -0.61359 | -0.16414 | -0.40797 | -0.53623 | -0.38007 | -0.28128 | -0.10488 | -0.006099 | 0.35609 | 0.31202 | -3.5802
Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20 | 0 | Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20 | 0
b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000
b2 | -2.0638 | -0.051021 | -0.4426 | 0.13584 | 0.70615 | 0.349 | 0.43686 | 1.0856 | 0.27648 | -0.21222 | 0.67978 | 0.23684 | 0.56453 | 0.92054 | 0.86058 | 0.31889 | 0.45938 | 0.88595 | 0.9089 | 0.14535 | 0.35813 | b2 | -4.7971 | 0.20419 | -0.14349 | -1.5695 | -0.49505 | -0.73748 | -0.16189 | -0.9544 | -0.79242 | -1.7941 | -1.4567 | -1.5843 | -1.0641 | -0.20739 | 0.051171 | -0.022432 | -1.5222 | -1.098 | 0.13367 | -0.72161 | -0.86781
b3 | 26.0355 | -2.1618 | -2.8936 | -2.7467 | -2.9981 | -3.0902 | -3.2382 | -3.1072 | -2.7766 | -3.188 | -3.0673 | -2.8117 | -2.6828 | -3.0028 | -2.8507 | -2.772 | -2.7103 | -2.8927 | -2.5156 | -5.2297 | -1.9861 | b3 | 28.936 | -2.0992 | -2.2896 | -2.0595 | -2.5694 | -2.7214 | -2.8192 | -2.3549 | -2.0073 | -2.9357 | -2.4187 | -2.4501 | -2.0415 | -1.9595 | -2.0864 | -2.3981 | -1.69 | -1.8963 | -1.5172 | -4.6058 | -1.2876
b4 | -1.3914 | 0.053072 | -0.029873 | -0.023156 | 0.012658 | -0.076898 | -0.049485 | 0.024348 | 0.086402 | -0.0026736 | -0.03115 | -0.038478 | -0.038093 | 0.042586 | -0.023974 | -0.057243 | 0.002847 | -0.030678 | 0.059237 | -0.29276 | -0.061722 | b4 | -0.74965 | -0.076509 | -0.010478 | -0.10808 | -0.057437 | -0.20661 | -0.23301 | -0.074624 | -0.064175 | -0.044832 | -0.21625 | -0.14461 | -0.070989 | -0.19913 | 0.010734 | -0.039225 | 0.076583 | -0.0092453 | 0.013693 | 0.012733 | -0.099805
b5 | -5.7754 | 0.14498 | 0.046533 | 0.060776 | 0.061861 | 0.34342 | 0.2787 | 0.25658 | 0.2691 | 0.32329 | 0.21441 | 0.1663 | 0.25166 | 0.22998 | 0.21043 | 0.15811 | 0.26815 | 0.15484 | 0.12417 | -4.1581 | -0.1417 | b5 | -5.7775 | 0.18639 | 0.20349 | 0.26972 | 0.35528 | 0.53981 | 0.19299 | 0.21942 | 0.27357 | 0.36161 | 0.25295 | 0.20006 | 0.2707 | 0.2466 | 0.32042 | 0.12679 | 0.25216 | 0.12273 | -0.071807 | -3.4449 | -0.079228
b6 | 0.50395 | 0.11768 | 0.056449 | -0.085771 | 0.090026 | 0.075011 | 0.082149 | -0.0097875 | 0.11242 | 0.15104 | -0.038485 | 0.10506 | -0.068075 | 0.01361 | -0.016817 | 0.094263 | -0.057394 | -0.018418 | 0.10032 | -0.0050193 | 0.052961 | b6 | -0.10598 | -0.13323 | -0.066877 | -0.067736 | -0.11701 | -0.11582 | -0.046592 | -0.2221 | -0.085561 | -0.025786 | -0.06414 | -0.01053 | -0.12075 | -0.075465 | -0.054078 | -0.0061507 | -0.10958 | -0.096624 | 0.055391 | -0.010278 | -0.074557
b7 | 1.6647 | 0.72354 | 0.82844 | 0.85758 | 0.82814 | 0.80357 | 0.87359 | 0.81673 | 0.83872 | 0.79825 | 0.82612 | 0.82137 | 0.86092 | 0.84236 | 0.77371 | 0.83755 | 0.8747 | 0.82469 | 0.82898 | 0.2716 | 0.83502 | b7 | 1.7078 | 0.7626 | 0.89842 | 0.87911 | 0.91606 | 0.85686 | 0.90335 | 0.87034 | 0.88759 | 0.8446 | 0.88488 | 0.89482 | 0.90973 | 0.9196 | 0.87183 | 0.8712 | 0.84397 | 0.82782 | 0.84553 | 0.17828 | 0.87648
b8 | 0.096574 | 0.064552 | 0.0291 | -0.14824 | 0.023236 | -0.037041 | 0.0062094 | -0.021231 | 0.042073 | 0.051612 | -0.085887 | 0.064947 | -0.071656 | 0.036849 | 0.024181 | 0.07202 | -0.032797 | 0.042885 | 0.067573 | -0.010899 | 0.0094948 | b8 | -0.071928 | 1.8514e-05 | -0.04836 | 0.016633 | -0.02818 | -0.0049359 | -0.028799 | -0.1479 | -0.039535 | 0.070896 | 0.0098976 | 0.034643 | -0.037551 | -0.009066 | 0.01907 | 0.019183 | 0.004589 | 0.0050925 | 0.042389 | -0.097212 | -0.011883
b9 | 0.44198 | 0.60334 | 0.57554 | 0.60148 | 0.59081 | 0.58578 | 0.58998 | 0.59568 | 0.59585 | 0.59012 | 0.59696 | 0.59051 | 0.58937 | 0.58617 | 0.58901 | 0.58765 | 0.59708 | 0.57843 | 0.59939 | 0.45432 | 0.58188 | b9 | 0.39137 | 0.60872 | 0.60379 | 0.59753 | 0.59915 | 0.58917 | 0.59139 | 0.59596 | 0.58717 | 0.6028 | 0.60419 | 0.60162 | 0.60765 | 0.58776 | 0.60951 | 0.59129 | 0.60857 | 0.60173 | 0.60415 | 0.4823 | 0.58885
b10 | 0.20918 | 0.068826 | 0.071918 | -0.2289 | 0.062297 | -0.039221 | 0.024303 | -0.074146 | 0.063381 | 0.12219 | -0.18217 | 0.080004 | -0.14693 | 0.038273 | -0.0041698 | 0.004281 | -0.11152 | 0.0091031 | 0.075186 | 0.063669 | 0.00030556 | b10 | 0.035514 | -0.027497 | -0.10609 | 0.023904 | -0.05224 | 0.029771 | -0.011741 | -0.20766 | -0.048031 | 0.10005 | 0.038202 | 0.059669 | -0.066174 | 0.015965 | 0.050363 | 0.069233 | 0.010831 | 0.046326 | 0.10153 | -0.054805 | 0.00075281
b11 | 0.60687 | 0.77097 | 0.76696 | 0.7653 | 0.76476 | 0.76557 | 0.76162 | 0.76515 | 0.76736 | 0.77387 | 0.77388 | 0.77111 | 0.76972 | 0.76857 | 0.77279 | 0.76546 | 0.76571 | 0.76303 | 0.76319 | 0.6318 | 0.75776 | b11 | 0.58659 | 0.76351 | 0.7648 | 0.76143 | 0.76575 | 0.76533 | 0.7586 | 0.76016 | 0.76202 | 0.76619 | 0.76064 | 0.76021 | 0.76196 | 0.7606 | 0.7685 | 0.76256 | 0.77065 | 0.76466 | 0.7655 | 0.65528 | 0.75412
b12 | 0.036897 | 0.0080949 | 0.008208 | -0.032985 | 0.0075716 | -0.0075935 | 0.0019986 | -0.0061723 | 0.0057549 | 0.011095 | -0.023297 | 0.011522 | -0.019408 | 0.0073324 | 0.0031759 | 0.0023788 | -0.013636 | 0.0031129 | 0.011591 | 0.0049555 | 0.00027651 | b12 | -0.0048772 | 0.0015772 | -0.011663 | 0.0043422 | -0.0053051 | 0.007839 | -0.00012234 | -0.028769 | -0.0069005 | 0.013545 | 0.0038607 | 0.0066638 | -0.0076473 | 0.0040446 | 0.0096194 | 0.011859 | 0.00020148 | 0.0041449 | 0.014043 | -0.014588 | 0.00046927
b13 | 0.083063 | 0.082549 | 0.085627 | 0.089809 | 0.086091 | 0.083808 | 0.084066 | 0.081457 | 0.083408 | 0.083469 | 0.087019 | 0.084998 | 0.086742 | 0.084035 | 0.084409 | 0.084345 | 0.086214 | 0.082064 | 0.083091 | 0.052939 | 0.083533 | b13 | 0.082833 | 0.084396 | 0.087665 | 0.084472 | 0.08491 | 0.081831 | 0.087518 | 0.086438 | 0.08648 | 0.08381 | 0.086276 | 0.087409 | 0.086614 | 0.086693 | 0.086861 | 0.087477 | 0.084453 | 0.085033 | 0.089793 | 0.047484 | 0.084641
b14 | 0.0020249 | -0.0014473 | -0.0019614 | -0.012649 | -0.0019154 | -0.0074148 | -0.0036171 | -0.0070546 | -0.0015036 | 0.0009605 | -0.012065 | -0.00061659 | -0.010079 | -0.0018511 | -0.0045711 | -0.0042419 | -0.009659 | -0.0045697 | -0.00033904 | -0.0037957 | -0.0044482 | b14 | 0.0024164 | -0.0055835 | -0.0078279 | -0.0031159 | -0.0073449 | -0.0016444 | -0.0033635 | -0.010685 | -0.0046585 | 0.0013511 | -0.0015408 | -0.00060898 | -0.0058143 | -0.0016622 | -0.0015738 | 0.00035681 | -0.0045269 | -0.0018462 | 0.0017865 | -0.0057776 | -0.0031352
b15 | 0.002472 | 0.030844 | 0.027607 | 0.01162 | 0.027521 | 0.026246 | 0.027923 | 0.023759 | 0.028244 | 0.032047 | 0.016833 | 0.029087 | 0.019373 | 0.026507 | 0.025469 | 0.028071 | 0.024626 | 0.031381 | 0.028365 | 0.014158 | 0.024708 | b15 | -0.00086662 | 0.030541 | 0.02344 | 0.032495 | 0.028181 | 0.033943 | 0.027796 | 0.013701 | 0.023664 | 0.029349 | 0.027297 | 0.028012 | 0.024263 | 0.028546 | 0.028999 | 0.028022 | 0.030721 | 0.030855 | 0.027076 | 0.014979 | 0.026252
b16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | b16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
b17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | b17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
a1 | 18.8241 | -8.2482 | -6.497 | -1.1477 | -2.5935 | 0.14966 | -1.4398 | -1.2653 | 0.90031 | 0.13652 | -5.1237 | -1.9961 | 2.163 | -6.7277 | -4.2955 | 4.9493 | 5.0844 | 5.4944 | 5.9499 | 4.3857 | 0 | a1 | 35.2246 | 3.6162 | 6.3676 | -2.4642 | -10.2557 | -2.276 | -1.8658 | 5.9754 | 5.3253 | -7.4684 | -6.1359 | -1.6414 | -4.0797 | -5.3623 | -3.8007 | -2.8128 | -1.0488 | -0.060989 | 3.561 | 3.1202 | -1.0386e-13
a2 | 9.5164 | -0.21749 | 0.26167 | -0.28232 | 0.11671 | 0.19313 | 0.84531 | 1.3531 | 1.0528 | 0.61795 | -0.094188 | 0.15876 | 0.61673 | 0.34182 | 0.14146 | -0.42609 | 0.48546 | 0.016364 | 0.22346 | -9.9728 | 0.26762 | a2 | 9.3715 | 1.3547 | 1.2407 | 1.4723 | 1.3681 | 1.0542 | 1.4179 | 2.0407 | 1.3041 | 1.1769 | 1.5194 | 1.3247 | 1.5278 | 1.3784 | 1.3027 | 1.5316 | 1.8122 | 1.541 | 2.1006 | -10.3166 | 1.3455
a3 | -3.0212 | -0.64295 | -0.85956 | -0.77732 | -0.84449 | -1.1531 | -1.3503 | -0.39723 | -0.50729 | -0.65813 | -0.23471 | -0.39993 | -0.43822 | -0.61129 | -0.64941 | -0.57639 | -1.0639 | -0.68675 | -0.50774 | 0.14103 | -0.73644 | a3 | -0.69418 | -0.17595 | -0.4763 | -0.52938 | -0.16154 | -0.47723 | -0.49568 | -0.35228 | -0.23583 | -0.52241 | 0.092318 | -0.042028 | 0.087491 | -0.044304 | -0.87355 | -1.0217 | -0.63041 | -0.61542 | -1.0455 | -0.88559 | -0.44223
a4 | 2.6065 | 0.20607 | 0.012266 | 0.21055 | 0.079676 | 0.38885 | 0.26607 | 0.3062 | 0.22162 | -0.029672 | -0.01609 | 0.1165 | 0.11908 | 0.030194 | 0.07755 | 0.12751 | 0.33749 | 0.055272 | 0.051359 | -0.15436 | 0.21378 | a4 | 1.5779 | 0.16592 | 0.085279 | 0.25612 | 0.63282 | 0.49678 | 0.046485 | 0.12301 | 0.22404 | 0.37541 | 0.42668 | 0.49564 | 0.41989 | 0.60952 | 0.34375 | 0.28211 | 0.26851 | 0.25218 | 0.12654 | -0.7786 | 0.31885
a5 | 1.3367 | -0.1318 | -0.18704 | -0.12701 | -0.15994 | -0.14913 | -0.25722 | -0.084677 | -0.028439 | -0.093211 | -0.13761 | -0.08878 | -0.12281 | -0.08068 | -0.13896 | -0.20392 | -0.17281 | -0.059998 | -0.11988 | -0.58608 | -0.096777 | a5 | 1.6208 | -0.092331 | -0.13483 | -0.17549 | -0.12986 | -0.032703 | -0.16782 | -0.1241 | -0.037571 | -0.17 | -0.087298 | -0.016137 | -0.024566 | 0.048223 | -0.19417 | -0.18492 | -0.14461 | -0.19763 | -0.29091 | -0.56584 | -0.076943
a6 | 0.033892 | 0.042485 | 0.089843 | 0.16762 | 0.093949 | 0.097656 | 0.10038 | 0.21261 | 0.13581 | 0.091547 | 0.15311 | 0.072923 | 0.17653 | 0.13125 | 0.083411 | 0.034913 | 0.046267 | 0.021004 | 0.049667 | -0.13684 | 0.090109 | a6 | -0.18941 | 0.024543 | -0.071728 | -0.087507 | 0.03095 | -0.025507 | -0.097404 | 0.0091237 | 0.011617 | -0.021017 | 0.0082439 | 0.0070975 | 0.07439 | 0.089158 | -0.046071 | -0.028279 | -0.088285 | -0.013156 | 0.041512 | 0.10453 | -0.012291
a7 | 1.7701 | 0.1263 | 0.071607 | 0.016504 | 0.024684 | 0.051631 | 0.044706 | 0.052753 | 0.10479 | 0.081492 | 0.017245 | 0.058248 | 0.047839 | 0.065118 | 0.064476 | 0.070608 | 0.045424 | 0.04808 | 0.052644 | 0.070396 | 0.11487 | a7 | 1.8741 | 0.073127 | 0.048956 | 0.049624 | 0.075129 | 0.16779 | 0.094429 | 0.063959 | 0.147 | 0.12426 | 0.095179 | 0.10988 | 0.13575 | 0.10763 | 0.028654 | 0.025213 | 0.060996 | 0.0644 | 0.03727 | 0.17106 | 0.14582
a8 | -0.046997 | -0.018795 | 0.021269 | 0.084949 | -0.014737 | -0.011417 | -0.027318 | -0.042814 | -0.032886 | -0.070276 | 0.038937 | -0.037639 | 0.070183 | 0.018161 | 0.016738 | -0.061717 | -0.042416 | -0.089627 | -0.062737 | -0.074966 | -0.017612 | a8 | -0.089879 | -0.025956 | 0.018581 | -0.032205 | 0.026751 | 0.00076902 | -0.017626 | 0.1217 | 0.067961 | 0.02269 | 0.040408 | -0.0028055 | 0.054542 | 0.025565 | 0.016873 | 0.0071511 | -0.0022602 | 0.022043 | 0.031192 | -0.049248 | 0.014901
a9 | -0.31141 | 0.05622 | 0.049165 | 0.0011294 | 0.040037 | 0.047633 | 0.04023 | 0.062387 | 0.046063 | 0.030133 | 0.05098 | 0.061108 | 0.046295 | 0.047965 | 0.038758 | 0.079508 | 0.081143 | 0.088852 | 0.05942 | 0.018806 | 0.038458 | a9 | -0.22032 | 0.082933 | 0.026725 | 0.041548 | 0.017758 | 0.056892 | 0.046715 | 0.024071 | 0.055418 | 0.015928 | 0.069863 | 0.042548 | 0.05396 | 0.033435 | 0.035114 | 0.056488 | 0.071083 | 0.067204 | 0.05109 | -0.040857 | 0.035442
a10 | -0.069892 | -0.042217 | 0.052305 | 0.11582 | -0.010043 | -0.0085898 | -0.023735 | -0.00086131 | -0.015264 | -0.064933 | 0.10853 | -0.046303 | 0.12437 | 0.053991 | 0.082337 | -0.021684 | -0.043931 | -0.10949 | -0.042115 | -0.10569 | -0.00015175 | a10 | -0.092078 | -0.020164 | 0.034508 | -0.079625 | -0.018544 | -0.056363 | -0.005284 | 0.17526 | 0.093067 | 0.00023885 | 0.014829 | -0.0135 | 0.03735 | -0.010655 | -0.0027446 | 0.014994 | -0.039842 | -0.041352 | 0.036689 | -0.08841 | 0.00023329
a11 | 0.25846 | 0.066763 | 0.060337 | 0.057093 | 0.060555 | 0.068528 | 0.065639 | 0.05988 | 0.073527 | 0.065503 | 0.050247 | 0.061879 | 0.060998 | 0.059727 | 0.06125 | 0.06947 | 0.070457 | 0.071832 | 0.06691 | 0.06824 | 0.070469 | a11 | 0.26454 | 0.067539 | 0.067071 | 0.058963 | 0.053588 | 0.077101 | 0.06355 | 0.063191 | 0.081118 | 0.063377 | 0.060181 | 0.068477 | 0.06519 | 0.066187 | 0.059033 | 0.065462 | 0.063169 | 0.06044 | 0.064455 | 0.092424 | 0.07239
a12 | -0.023779 | -0.0044778 | 0.007767 | 0.017827 | 0.001096 | 0.0015517 | 0.00052619 | 0.0071679 | 0.0023061 | -0.0053362 | 0.016527 | -0.0048377 | 0.019278 | 0.0093526 | 0.010716 | -0.0021378 | -0.0030559 | -0.010951 | -0.003011 | -0.011596 | 0.0019662 | a12 | -0.017591 | -0.0034266 | -0.0015339 | -0.014358 | -0.0028037 | -0.0077385 | -0.0057616 | 0.017256 | 0.010541 | -0.0022136 | 0.00088651 | -0.00053896 | 0.0074294 | -0.00023425 | -0.003 | -0.001368 | -0.0077697 | -0.0076488 | 0.0036634 | -0.0038417 | -0.001673
a13 | -0.018959 | 0.0079825 | 0.0063781 | 0.0050896 | 0.0049767 | 0.0065018 | 0.0037638 | 0.0082212 | 0.0070324 | 0.0061333 | 0.0080651 | 0.0085992 | 0.0084025 | 0.0057603 | 0.0052233 | 0.0096693 | 0.010093 | 0.010322 | 0.0081698 | 0.0037508 | 0.006264 | a13 | -0.010616 | 0.012525 | 0.0080936 | 0.007268 | 0.004073 | 0.0078196 | 0.0062694 | 0.0077304 | 0.0094499 | 0.0043428 | 0.0086339 | 0.0087027 | 0.010323 | 0.0083164 | 0.0055749 | 0.0074085 | 0.0084807 | 0.0075356 | 0.0048276 | -0.0012332 | 0.0067394
a14 | 0.011186 | -0.0097026 | -0.0058373 | -0.0019979 | -0.0076412 | -0.0061844 | -0.0071838 | -0.0075132 | -0.0078995 | -0.012139 | -0.0028801 | -0.010197 | -0.0015907 | -0.0053737 | -0.0050282 | -0.010211 | -0.0070264 | -0.013958 | -0.010224 | -0.004007 | -0.0066443 | a14 | 0.0068378 | -0.011256 | -0.0071458 | -0.012836 | -0.0080703 | -0.010698 | -0.0094664 | -0.0019365 | -0.0043676 | -0.0076951 | -0.0065383 | -0.0085739 | -0.0061221 | -0.0084756 | -0.0084241 | -0.0084926 | -0.010596 | -0.010356 | -0.0072692 | -0.0039953 | -0.0076032
a15 | 0.0096899 | 0.0011188 | 0.0017969 | -5.4093e-05 | -9.2627e-05 | -0.0042947 | -0.0030199 | -0.0028188 | 0.0014178 | 0.0029712 | 0.00060038 | 0.0014137 | 0.0016765 | 0.0011644 | 0.0008212 | 0.00010219 | -0.0075168 | -0.0027869 | 0.0017949 | -0.0050318 | -0.00014162 | a15 | -0.0033195 | -0.0025908 | -0.0043197 | -0.0018335 | -0.005266 | -0.00075535 | -0.0016213 | 0.0028192 | 0.000607 | 0.0044577 | 0.0008156 | 0.002037 | -0.0016777 | -3.7486e-05 | -0.00042475 | 0.00030498 | -0.0013647 | 1.9787e-05 | 0.0013497 | -0.015458 | -0.0010039
a16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | a16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
a17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | a17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Dx (mm) | 0.42543 | 0.14263 | 0.16749 | -0.48656 | 0.13586 | -0.088094 | 0.049315 | -0.16388 | 0.13592 | 0.25454 | -0.38309 | 0.16711 | -0.30085 | 0.093371 | 0.0051507 | 0.005095 | -0.25442 | -0.0026588 | 0.15804 | 0.13898 |  | Dx (mm) | -0.014476 | -0.064688 | -0.22734 | 0.039367 | -0.11828 | 0.052979 | -0.027114 | -0.42887 | -0.084098 | 0.22051 | 0.087456 | 0.12966 | -0.13949 | 0.033357 | 0.11029 | 0.15606 | 0.016577 | 0.095133 | 0.2307 | -0.17113
Dy (mm) | -0.37697 | -0.10544 | 0.10276 | 0.29357 | -0.033082 | -0.011189 | -0.057229 | 0.010912 | -0.04684 | -0.16419 | 0.26327 | -0.11549 | 0.29854 | 0.11217 | 0.18072 | -0.04862 | -0.074125 | -0.24368 | -0.10767 | -0.29938 |  | Dy (mm) | -0.26032 | -0.039175 | 0.096641 | -0.18082 | -0.032892 | -0.13053 | -0.0095699 | 0.4276 | 0.21658 | -0.01771 | 0.026222 | -0.041868 | 0.095266 | -0.026718 | -0.014544 | 0.020028 | -0.089248 | -0.099453 | 0.062053 | -0.20522
Mag. len. (m) | 14.4479 | [20 empty] | Mag. len. (m) | 14.4484
TF (mT/KA) | 595.0719 | [20 empty] | TF (mT/KA) | 595.0196
Coil pos. (mm) | 0.013299 | [20 empty] | Coil pos. (mm) | 0.016155
Current (A) | 8.5 | [20 empty] | Current (A) | 8.5

## Harmonics
Normal:
Skew:
*10

## Work sheet
| Measured harmonics- Aperture 1 | [7 empty] | Measured harmonics- Aperture 2 | [8 empty] | Zero vector
| normal multipoles |  |  |  | skew multipoles |  |  |  | normal multipoles |  |  |  | skew multipoles
| all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19)
| average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma
1 | [17 empty] | 0
2 | 0.35813 | 0.68958 | 0.44404 | 0.41025 | 0.26762 | 3.1962 | 0.30409 | 0.46282 | -0.86781 | 1.1106 | -0.73331 | 0.66749 | 1.3455 | 3.2597 | 1.4769 | 0.29378 |  | 0
3 | -1.9861 | 6.5155 | -2.862 | 0.25835 | -0.73644 | 0.62421 | -0.68444 | 0.28395 | -1.2876 | 7.0281 | -2.2403 | 0.38135 | -0.44223 | 0.34886 | -0.41918 | 0.34971 |  | 0
4 | -0.061722 | 0.3157 | -0.0062224 | 0.043165 | 0.21378 | 0.57083 | 0.14111 | 0.12458 | -0.099805 | 0.17349 | -0.080612 | 0.089599 | 0.31885 | 0.41737 | 0.31378 | 0.1785 |  | 0
5 | -0.1417 | 1.6134 | 0.19642 | 0.088686 | -0.096777 | 0.3524 | -0.12603 | 0.055847 | -0.079228 | 1.5449 | 0.23861 | 0.12494 | -0.076943 | 0.41442 | -0.11853 | 0.084557 |  | 0
6 | 0.052961 | 0.12615 | 0.045665 | 0.036651 | 0.090109 | 0.07417 | 0.094712 | 0.040464 | -0.074557 | 0.059496 | -0.076211 | 0.038253 | -0.012291 | 0.070564 | -0.014941 | 0.056655 |  | 0
7 | 0.83502 | 0.23272 | 0.82357 | 0.035705 | 0.11487 | 0.3836 | 0.063004 | 0.025657 | 0.87648 | 0.25169 | 0.87057 | 0.038374 | 0.14582 | 0.40169 | 0.086875 | 0.042607 |  | 0
8 | 0.0094948 | 0.061695 | 0.012561 | 0.023299 | -0.017612 | 0.04783 | -0.018373 | 0.015115 | -0.011883 | 0.050514 | -0.0080967 | 0.014449 | 0.014901 | 0.044609 | 0.016402 | 0.017931 |  | 0
9 | 0.58188 | 0.044645 | 0.58585 | 0.0095349 | 0.038458 | 0.08329 | 0.061623 | 0.018912 | 0.58885 | 0.052697 | 0.59725 | 0.0075903 | 0.035442 | 0.064575 | 0.053529 | 0.023365 |  | 0
10 | 0.00030556 | 0.10782 | 1.2784e-09 | 2.1137e-08 | -0.00015175 | 0.07041 | -7.0575e-09 | 1.4911e-08 | 0.00075281 | 0.074549 | -3.2641e-09 | 2.0205e-08 | 0.00023329 | 0.062305 | -1.1728e-09 | 7.6591e-09 |  | 0
11 | 0.75776 | 0.045922 | 0.76666 | 0.0041502 | 0.070469 | 0.043812 | 0.065611 | 0.0048655 | 0.75412 | 0.045373 | 0.76305 | 0.0033609 | 0.07239 | 0.045146 | 0.065887 | 0.0071205 |  | 0
12 | 0.00027651 | 0.015132 | 0.00021725 | 0.002182 | 0.0019662 | 0.010686 | 0.0024998 | 0.0017426 | 0.00046927 | 0.010573 | 0.0010049 | 0.0020175 | -0.001673 | 0.0078755 | -0.0019402 | 0.0024666 |  | 0
13 | 0.083533 | 0.0073322 | 0.084198 | 0.0012942 | 0.006264 | 0.0061349 | 0.0078515 | 0.0017826 | 0.084641 | 0.0087796 | 0.085784 | 0.0017553 | 0.0067394 | 0.0048696 | 0.0080183 | 0.0020859 |  | 0
14 | -0.0044482 | 0.0042322 | -0.0049853 | 0.0010849 | -0.0066443 | 0.0052487 | -0.0072902 | 0.0011495 | -0.0031352 | 0.0034483 | -0.0036777 | 0.00069701 | -0.0076032 | 0.004226 | -0.0082135 | 0.0010559 |  | 0
15 | 0.024708 | 0.0075619 | 0.025862 | 0.005224 | -0.00014162 | 0.0035887 | -0.000317 | 0.0026979 | 0.026252 | 0.0080228 | 0.027606 | 0.0044544 | -0.0010039 | 0.0040713 | -0.00041557 | 0.0023804 |  | 0
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0
Coil positioning estimate
1 | [12 empty] | n | n^2 | Ap. 1 | Ap. 2
2 | [12 empty] | 2 | 4 | -0.8307 | -0.81457
3 | [12 empty] | 3 | 9 | -1.3062 | -1.0073
4 | [12 empty] | 4 | 16 | -2.6128 | -2.0678
5 | [12 empty] | 5 | 25 | -2.6539 | -2.2751
6 | [12 empty] | 6 | 36 | -3.2568 | -3.0672
7 | [12 empty] | 7 | 49 | -3.4977 | -3.2081
8 | [12 empty] | 8 | 64 | -3.9757 | -4.1292
9 | [12 empty] | 9 | 81 | -4.3104 | -4.3187
10 | [12 empty] | sigma (mm) |  | 0.013299 | 0.016155
11
12
13
14
15
16
17
| Random harmonics | [6 empty] | Expected harmonics (no uncertainty)
| d (mm) |  |  |  | sig(n)=d alpha beta^n |  |  | normal |  |  | skew
| 0.1 | 0.025 | 0.006 |  |  |  |  | min | ave | max | min | ave | max
1 | 4.6023 | 1.1506 | 0.27614 |  | Scaling law constants |  |  | 0 | 0 | 0 | 0 | 0 | 0
2 | 2.8314 | 0.70784 | 0.16988 |  | alpha | 73.9 |  | 0 | 0 | 0 | 0 | 0 | 0
3 | 1.7207 | 0.43017 | 0.10324 |  | beta | 0.6266 |  | 0 | 0 | 0 | 0 | 0 | 0
4 | 1.033 | 0.25824 | 0.061978 |  | gamma | 0.9939 |  | 0 | 0 | 0 | 0 | 0 | 0
5 | 0.61258 | 0.15315 | 0.036755 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
6 | 0.35886 | 0.089715 | 0.021532 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
7 | 0.20767 | 0.051917 | 0.01246 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
8 | 0.11871 | 0.029678 | 0.0071228 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
9 | 0.067037 | 0.016759 | 0.0040222 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
10 | 0.037395 | 0.0093488 | 0.0022437 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
11 | 0.020607 | 0.0051516 | 0.0012364 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
12 | 0.011217 | 0.0028042 | 0.00067302 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
13 | 0.0060316 | 0.0015079 | 0.0003619 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
14 | 0.0032039 | 0.00080097 | 0.00019223 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
15 | 0.0016811 | 0.00042029 | 0.00010087 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
16 | 0.0008714 | 0.00021785 | 5.2284e-05 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
17 | 0.00044619 | 0.00011155 | 2.6771e-05 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
Measured harmonics along the axis with local feed down - Aperture 1 - Normal
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20
2 | -3.2329 | -0.0067707 | -0.39597 | -0.048237 | 0.75736 | 0.31849 | 0.46474 | 1.0252 | 0.32367 | -0.10405 | 0.53427 | 0.29756 | 0.45418 | 0.94546 | 0.8485 | 0.32385 | 0.38753 | 0.90474 | 0.9621 | 0.22589 | 0.35813
3 | 25.9665 | -2.1669 | -2.8925 | -2.7378 | -2.9989 | -3.0922 | -3.2405 | -3.1059 | -2.7805 | -3.187 | -3.0702 | -2.8129 | -2.6785 | -3.0029 | -2.8482 | -2.7731 | -2.7145 | -2.8951 | -2.5182 | -5.2144 | -1.9861
4 | -0.93183 | 0.051476 | -0.03623 | -0.024971 | 0.011925 | -0.069387 | -0.049256 | 0.034024 | 0.078109 | -0.018435 | -0.020348 | -0.042605 | -0.028905 | 0.035405 | -0.030138 | -0.0551 | 0.021913 | -0.027141 | 0.057657 | -0.1155 | -0.061722
5 | -5.8422 | 0.13873 | 0.046468 | 0.062974 | 0.05735 | 0.34504 | 0.27582 | 0.25679 | 0.26273 | 0.30756 | 0.22193 | 0.15866 | 0.26113 | 0.23393 | 0.21489 | 0.15747 | 0.26284 | 0.15332 | 0.11794 | -4.1458 | -0.1417
6 | 0.018485 | 0.076557 | 0.010073 | 0.063209 | 0.050027 | 0.099792 | 0.066041 | 0.037656 | 0.070453 | 0.074608 | 0.074815 | 0.054236 | 0.028381 | -0.011572 | -0.014111 | 0.091545 | 0.019963 | -0.021779 | 0.052076 | -0.02578 | 0.052961
7 | 1.6551 | 0.72057 | 0.82733 | 0.83815 | 0.82705 | 0.80228 | 0.87411 | 0.81511 | 0.837 | 0.79759 | 0.8168 | 0.81869 | 0.86067 | 0.84178 | 0.77491 | 0.83863 | 0.87256 | 0.83373 | 0.82736 | 0.28146 | 0.83502
8 | 0.063333 | 0.021266 | -0.013887 | -0.010366 | -0.01516 | -0.013008 | -0.0085656 | 0.025029 | 0.002945 | -0.021401 | 0.028047 | 0.015188 | 0.01829 | 0.013624 | 0.026049 | 0.068792 | 0.035859 | 0.03342 | 0.019986 | -0.043262 | 0.0094948
9 | 0.40881 | 0.6005 | 0.57201 | 0.56052 | 0.5865 | 0.584 | 0.59006 | 0.58924 | 0.59167 | 0.57929 | 0.57514 | 0.58626 | 0.58563 | 0.58749 | 0.5969 | 0.5882 | 0.58378 | 0.59257 | 0.5955 | 0.46639 | 0.58188
10 | 1.2516e-08 | -1.6044e-08 | 1.0308e-08 | -5.4221e-08 | 6.9185e-09 | -1.2407e-09 | -2.0274e-09 | -2.0231e-08 | 1.2284e-08 | -1.4906e-08 | 5.8329e-09 | 6.4001e-09 | 4.9774e-08 | -1.5031e-09 | -1.4579e-09 | 4.2262e-10 | 1.8351e-08 | 3.6318e-11 | 2.4314e-08 | -1.6772e-08 | 0.00030556
11 | 0.60252 | 0.77053 | 0.76659 | 0.7583 | 0.76407 | 0.76512 | 0.76154 | 0.76455 | 0.76679 | 0.77261 | 0.77092 | 0.77022 | 0.76966 | 0.76881 | 0.77403 | 0.76552 | 0.76361 | 0.76477 | 0.76222 | 0.6336 | 0.75776
12 | 0.16998 | -0.0081038 | -0.01453 | -0.010845 | -0.0080095 | -0.024333 | -0.010798 | 0.033141 | -0.024802 | -0.046128 | 0.017333 | 0.0079435 | 0.0078355 | 0.022498 | 0.035353 | 0.017436 | 0.013193 | 0.014914 | 0.017006 | -0.010306 | 0.0027651 | ten times!
13 | 0.7918 | 0.83489 | 0.85419 | 0.84654 | 0.86483 | 0.83362 | 0.84516 | 0.8051 | 0.83847 | 0.84806 | 0.82905 | 0.85978 | 0.8406 | 0.83706 | 0.83733 | 0.84741 | 0.84733 | 0.84656 | 0.83973 | 0.5426 | 0.83533 | ten times!
14 | -0.018494 | -0.051674 | -0.056173 | -0.080065 | -0.049921 | -0.054712 | -0.046088 | -0.038734 | -0.047198 | -0.061589 | -0.066241 | -0.04754 | -0.048668 | -0.037818 | -0.045569 | -0.043638 | -0.040405 | -0.039417 | -0.041899 | -0.041755 | -0.044482 | ten times!
15 | 0.02472 | 0.30844 | 0.27607 | 0.1162 | 0.27521 | 0.26246 | 0.27923 | 0.23759 | 0.28244 | 0.32047 | 0.16833 | 0.29087 | 0.19373 | 0.26507 | 0.25469 | 0.28071 | 0.24626 | 0.31381 | 0.28365 | 0.14158 | 0.24708 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Measured harmonics along the axis with local feed down- Aperture 1 - Skew
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1
2 | 10.8223 | -0.23352 | 0.31359 | -0.23195 | 0.11854 | 0.17711 | 0.83134 | 1.3494 | 1.0456 | 0.57608 | -0.0097611 | 0.12842 | 0.69545 | 0.38816 | 0.20246 | -0.4416 | 0.42999 | -0.066779 | 0.20103 | -10.1593 | 0.26762
3 | -3.3094 | -0.64715 | -0.85938 | -0.75805 | -0.84633 | -1.1472 | -1.3531 | -0.38843 | -0.51189 | -0.65688 | -0.23435 | -0.40415 | -0.42989 | -0.61263 | -0.64872 | -0.57699 | -1.0487 | -0.68804 | -0.50805 | 0.12935 | -0.73644
4 | 1.9604 | 0.21409 | 0.018512 | 0.19181 | 0.085271 | 0.38666 | 0.2728 | 0.30227 | 0.2255 | -0.0116 | -0.041776 | 0.12451 | 0.092706 | 0.025897 | 0.06877 | 0.12956 | 0.33183 | 0.064112 | 0.058963 | -0.4281 | 0.21378
5 | 1.3883 | -0.12993 | -0.19317 | -0.095618 | -0.16281 | -0.14635 | -0.2573 | -0.074397 | -0.032319 | -0.092771 | -0.11738 | -0.088796 | -0.10121 | -0.084733 | -0.13819 | -0.20262 | -0.1706 | -0.061302 | -0.11901 | -0.58093 | -0.096777
6 | -0.010415 | 0.063053 | 0.055564 | 0.081593 | 0.10243 | 0.10243 | 0.11725 | 0.21252 | 0.14464 | 0.13048 | 0.078675 | 0.10297 | 0.090896 | 0.095756 | 0.033944 | 0.049158 | 0.073229 | 0.091976 | 0.078232 | -0.1116 | 0.090109
7 | 1.7934 | 0.1302 | 0.068909 | 0.051443 | 0.025825 | 0.051046 | 0.045407 | 0.04996 | 0.10745 | 0.092347 | 0.032698 | 0.063926 | 0.065342 | 0.062718 | 0.062641 | 0.07218 | 0.039979 | 0.052285 | 0.059722 | 0.073342 | 0.11487
8 | 0.093754 | 0.0073691 | -0.010438 | 0.0021113 | -0.0080989 | -0.0063581 | -0.012363 | -0.041062 | -0.022698 | -0.02829 | -0.025832 | -0.010351 | -0.0060619 | -0.014887 | -0.033447 | -0.048463 | -0.011874 | -0.023185 | -0.036787 | -0.012188 | -0.017612
9 | -0.25392 | 0.06325 | 0.040615 | 0.066539 | 0.04185 | 0.047 | 0.041586 | 0.062741 | 0.048733 | 0.049505 | 0.098381 | 0.070096 | 0.089327 | 0.043024 | 0.038933 | 0.079677 | 0.07085 | 0.089872 | 0.06723 | 0.036673 | 0.038458
10 | -8.7812e-10 | -1.0508e-08 | -2.5274e-08 | -4.4984e-08 | 2.6317e-09 | -2.4525e-09 | 2.2084e-09 | -2.0475e-09 | -3.8088e-09 | 1.7816e-09 | -1.021e-08 | -8.3847e-09 | -3.3416e-08 | 1.3696e-08 | 1.2079e-08 | 8.7671e-10 | 8.8221e-10 | -1.4261e-08 | -5.8442e-09 | -2.1511e-08 | -0.00015175
11 | 0.27401 | 0.067728 | 0.058949 | 0.068971 | 0.060621 | 0.068562 | 0.065697 | 0.060683 | 0.073499 | 0.06756 | 0.058312 | 0.063263 | 0.0685 | 0.05863 | 0.060843 | 0.069552 | 0.0693 | 0.072304 | 0.068026 | 0.070243 | 0.070469 | ten times!
12 | 0.04017 | 0.0086249 | 0.0080181 | 0.0096446 | 0.026291 | 0.026179 | 0.037912 | 0.074915 | 0.043892 | 0.032355 | 0.02536 | 0.010773 | 0.027834 | 0.023194 | -0.00070555 | 0.0072212 | 0.032677 | 0.031846 | 0.023925 | -0.0077613 | 0.019662 | ten times!
13 | -0.22014 | 0.089241 | 0.072799 | 0.07186 | 0.057221 | 0.060218 | 0.038764 | 0.073385 | 0.077996 | 0.086168 | 0.096504 | 0.098478 | 0.10338 | 0.063028 | 0.058748 | 0.095514 | 0.081781 | 0.094424 | 0.093775 | 0.033077 | 0.06264 | ten times!
14 | 0.085585 | -0.071558 | -0.084215 | -0.048289 | -0.06881 | -0.062541 | -0.057451 | -0.081071 | -0.069687 | -0.08429 | -0.063403 | -0.076251 | -0.059384 | -0.079117 | -0.088221 | -0.090879 | -0.070982 | -0.076662 | -0.079427 | 0.00059445 | -0.066443 | ten times!
15 | 0.096899 | 0.011188 | 0.017969 | -0.00054093 | -0.00092627 | -0.042947 | -0.030199 | -0.028188 | 0.014178 | 0.029712 | 0.0060038 | 0.014137 | 0.016765 | 0.011644 | 0.008212 | 0.0010219 | -0.075168 | -0.027869 | 0.017949 | -0.050318 | -0.0014162 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Measured harmonics along the axis with local feed down - Aperture 2 - Normal
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1
2 | -4.7266 | 0.18903 | -0.21015 | -1.5487 | -0.53017 | -0.71318 | -0.17033 | -1.0909 | -0.81829 | -1.7168 | -1.4316 | -1.5467 | -1.0966 | -0.19956 | 0.079738 | 0.019191 | -1.5123 | -1.0695 | 0.16722 | -0.79295 | -0.86781
3 | 28.8616 | -2.1012 | -2.2886 | -2.0669 | -2.5743 | -2.7309 | -2.8204 | -2.3512 | -1.9997 | -2.9352 | -2.4134 | -2.4504 | -2.0362 | -1.9612 | -2.0875 | -2.396 | -1.6945 | -1.9006 | -1.5164 | -4.5772 | -1.2876
4 | -0.86861 | -0.07282 | -0.0026584 | -0.10311 | -0.046545 | -0.21234 | -0.2314 | -0.064968 | -0.060676 | -0.062886 | -0.22199 | -0.15055 | -0.062655 | -0.20137 | 0.003083 | -0.044752 | 0.078636 | -0.007368 | 0.013344 | -0.098652 | -0.099805
5 | -5.7635 | 0.18357 | 0.19698 | 0.27516 | 0.35091 | 0.54259 | 0.19289 | 0.19255 | 0.27219 | 0.36339 | 0.25466 | 0.20038 | 0.26783 | 0.24664 | 0.32238 | 0.1269 | 0.25501 | 0.12582 | -0.074808 | -3.4517 | -0.079228
6 | -0.26944 | -0.11683 | 0.0068783 | -0.083117 | -0.079638 | -0.13958 | -0.038266 | -0.080708 | -0.047979 | -0.092297 | -0.090572 | -0.053103 | -0.071396 | -0.087306 | -0.088163 | -0.053959 | -0.11644 | -0.12668 | -0.012639 | -0.011901 | -0.074557
7 | 1.717 | 0.76302 | 0.89463 | 0.88124 | 0.91433 | 0.85693 | 0.9031 | 0.86565 | 0.89228 | 0.838 | 0.88496 | 0.89302 | 0.90971 | 0.91945 | 0.87086 | 0.87002 | 0.84402 | 0.82672 | 0.8423 | 0.1756 | 0.87648
8 | -0.042272 | 0.01702 | 0.01745 | 0.0020281 | 0.0048942 | -0.023119 | -0.021464 | -0.022782 | -0.01065 | 0.00821 | -0.014106 | -0.0029049 | 0.0047552 | -0.018713 | -0.012805 | -0.023711 | -0.0031438 | -0.024991 | -0.021709 | -0.054427 | -0.011883
9 | 0.40433 | 0.6082 | 0.59279 | 0.60465 | 0.5962 | 0.59223 | 0.59125 | 0.58848 | 0.5957 | 0.59111 | 0.60263 | 0.59782 | 0.60465 | 0.58763 | 0.6066 | 0.58573 | 0.61036 | 0.60157 | 0.59296 | 0.48694 | 0.58885
10 | -1.9842e-09 | 7.0635e-09 | -2.3181e-08 | -3.3967e-09 | 3.2997e-09 | 7.3975e-09 | -3.7711e-09 | 4.3076e-08 | 7.5294e-11 | -4.5593e-08 | 1.2458e-09 | 7.8505e-09 | 4.3661e-09 | 5.5444e-09 | -1.2182e-08 | -2.2269e-09 | -4.4861e-10 | -4.1407e-10 | -4.7459e-08 | 1.69e-08 | 0.00075281
11 | 0.58951 | 0.76367 | 0.76299 | 0.763 | 0.7654 | 0.76572 | 0.75863 | 0.75695 | 0.76313 | 0.76429 | 0.76043 | 0.75967 | 0.76173 | 0.76052 | 0.76784 | 0.76135 | 0.7711 | 0.76489 | 0.76355 | 0.65418 | 0.75412
12 | -0.020801 | 0.050846 | 0.029571 | 0.010672 | 0.016895 | 0.040583 | 0.015103 | -0.0026816 | -0.0032214 | 0.0044519 | -0.013056 | -0.015936 | 0.015752 | 0.018465 | 0.027998 | 0.023276 | -0.01321 | -0.020943 | -0.0036771 | -0.086734 | 0.0046927 | ten times!
13 | 0.81499 | 0.84457 | 0.85777 | 0.86341 | 0.84449 | 0.82966 | 0.87517 | 0.823 | 0.85457 | 0.83687 | 0.86248 | 0.87744 | 0.85548 | 0.86909 | 0.87088 | 0.87304 | 0.85234 | 0.85955 | 0.89133 | 0.47354 | 0.84641 | ten times!
14 | 0.031177 | -0.038729 | -0.037833 | -0.038964 | -0.044573 | -0.030441 | -0.027301 | -0.04853 | -0.029113 | -0.040436 | -0.034891 | -0.036704 | -0.031587 | -0.024455 | -0.042026 | -0.032396 | -0.04846 | -0.042651 | -0.032888 | -0.010541 | -0.031352 | ten times!
15 | -0.0086662 | 0.30541 | 0.2344 | 0.32495 | 0.28181 | 0.33943 | 0.27796 | 0.13701 | 0.23664 | 0.29349 | 0.27297 | 0.28012 | 0.24263 | 0.28546 | 0.28999 | 0.28022 | 0.30721 | 0.30855 | 0.27076 | 0.14979 | 0.26252 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Measured harmonics along the axis with local feed down - Aperture 2 - Skew
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1
2 | 10.2565 | 1.3437 | 1.254 | 1.431 | 1.3559 | 1.0154 | 1.4131 | 2.1414 | 1.3529 | 1.1844 | 1.5259 | 1.3133 | 1.5521 | 1.3724 | 1.3104 | 1.556 | 1.7957 | 1.5257 | 2.1401 | -10.4456 | 1.3455
3 | -0.72459 | -0.17459 | -0.4727 | -0.53461 | -0.14866 | -0.48663 | -0.49585 | -0.33734 | -0.23005 | -0.53715 | 0.086734 | -0.054438 | 0.09902 | -0.048831 | -0.88021 | -1.0293 | -0.62999 | -0.61982 | -1.0508 | -0.90864 | -0.44223
4 | 1.2295 | 0.16623 | 0.07344 | 0.26922 | 0.63196 | 0.51376 | 0.045849 | 0.088409 | 0.20935 | 0.38573 | 0.42692 | 0.4981 | 0.41302 | 0.61069 | 0.34989 | 0.2883 | 0.27436 | 0.25948 | 0.14338 | -0.96773 | 0.31885
5 | 1.6119 | -0.093399 | -0.13773 | -0.17808 | -0.12991 | -0.036752 | -0.16872 | -0.095013 | -0.031834 | -0.16877 | -0.087015 | -0.016538 | -0.01813 | 0.046755 | -0.1929 | -0.18359 | -0.14706 | -0.20009 | -0.29474 | -0.5612 | -0.076943
6 | -0.02292 | 0.036756 | -0.098444 | -0.032093 | 0.044721 | 0.010832 | -0.093449 | -0.11255 | -0.051865 | -0.025408 | -0.0028834 | 0.015292 | 0.050485 | 0.096562 | -0.042711 | -0.035826 | -0.062057 | 0.01374 | 0.019959 | 0.12778 | -0.012291
7 | 1.8658 | 0.072436 | 0.052619 | 0.051384 | 0.076051 | 0.16751 | 0.094119 | 0.11149 | 0.15288 | 0.12272 | 0.093617 | 0.11063 | 0.14036 | 0.10718 | 0.028002 | 0.024596 | 0.06118 | 0.063745 | 0.033223 | 0.15938 | 0.14582
8 | -0.043436 | -0.01221 | -0.0060191 | 0.017871 | 0.037013 | 0.035542 | -0.014367 | 0.0066339 | 0.010312 | 0.026061 | 0.030077 | 0.006452 | 0.030843 | 0.03243 | 0.019222 | -0.0025705 | 0.022745 | 0.047196 | 0.0080039 | -0.0059601 | 0.014901
9 | -0.21613 | 0.081672 | 0.036306 | 0.045496 | 0.015687 | 0.06053 | 0.046579 | 0.11087 | 0.065069 | 0.016838 | 0.068646 | 0.044797 | 0.060056 | 0.033849 | 0.035662 | 0.054516 | 0.071944 | 0.071726 | 0.043274 | -0.054821 | 0.035442
10 | 1.1494e-08 | 1.834e-09 | -4.0613e-09 | 1.3996e-08 | 2.5287e-09 | -1.9717e-08 | 3.0393e-10 | 1.3016e-08 | -1.0973e-08 | -6.5552e-10 | 9.1629e-10 | -5.6038e-09 | 3.698e-09 | -1.8367e-09 | -1.1919e-09 | -4.9055e-10 | -2.5682e-09 | -7.3992e-09 | -2.9052e-09 | -9.3718e-09 | 0.00023329
11 | 0.26356 | 0.067435 | 0.067575 | 0.059836 | 0.05326 | 0.078028 | 0.063448 | 0.075939 | 0.082659 | 0.063848 | 0.060065 | 0.068702 | 0.066332 | 0.066262 | 0.059338 | 0.065447 | 0.063263 | 0.061178 | 0.063344 | 0.090061 | 0.07239
12 | -0.024785 | -0.0052082 | -0.062154 | -0.037777 | -0.0049214 | -0.0049087 | -0.050504 | -0.064939 | -0.021188 | -0.018419 | -0.012434 | 0.012478 | 0.026213 | 0.012049 | -0.025423 | -0.034209 | -0.025485 | -0.021853 | -0.010558 | 0.02888 | -0.01673 | ten times!
13 | -0.10059 | 0.11801 | 0.074298 | 0.072236 | 0.031583 | 0.080889 | 0.060485 | 0.10589 | 0.099406 | 0.056587 | 0.09102 | 0.095334 | 0.10093 | 0.084986 | 0.062679 | 0.084166 | 0.083061 | 0.081486 | 0.060252 | -0.026628 | 0.067394 | ten times!
14 | 0.066124 | -0.10409 | -0.098201 | -0.079381 | -0.078198 | -0.070161 | -0.092835 | -0.057654 | -0.085463 | -0.080765 | -0.071865 | -0.078255 | -0.082184 | -0.078464 | -0.080382 | -0.08994 | -0.083197 | -0.078306 | -0.089092 | -0.036421 | -0.076032 | ten times!
15 | -0.033195 | -0.025908 | -0.043197 | -0.018335 | -0.05266 | -0.0075535 | -0.016213 | 0.028192 | 0.00607 | 0.044577 | 0.008156 | 0.02037 | -0.016777 | -0.00037486 | -0.0042475 | 0.0030498 | -0.013647 | 0.00019787 | 0.013497 | -0.15458 | -0.010039 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Relative main field module along the axis
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
|  | 0.00057114 | -2.1967e-05 | -2.1967e-05 | -0.00021967 | -0.00021967 | -0.00021967 | -0.00021967 | -2.1967e-05 | -2.1967e-05 | -2.1967e-05 | -2.1967e-05 | -2.1967e-05 | -2.1967e-05 | 0.00017574 | -2.1967e-05 | -0.00021967 | -2.1967e-05 | 0.00057114 |  | 595.0719
|  | 0.00065907 | 6.5907e-05 | -0.00013181 | -0.00013181 | -0.00013181 | -0.00032953 | -0.00032953 | 6.5907e-05 | -0.00013181 | -0.00013181 | -0.00013181 | -0.00013181 | 6.5907e-05 | 6.5907e-05 | -0.00013181 | 6.5907e-05 | 6.5907e-05 | 0.00065907 |  | 595.0196
Check Computation for Aperture1
Magnetic Length
mbh dB/B | -3668.1331
| mean 3:18 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
dB/B | 595.3805 | -3647.3728 | 6.4258 | 0.49429 | 0.49429 | -1.4829 | -1.4829 | -1.4829 | -1.4829 | 0.49429 | 0.49429 | 0.49429 | 0.49429 | 0.49429 | 0.49429 | 2.4715 | 0.49429 | -1.4829 | 0.49429 | 6.4258 | -3688.8933
| mean 2:19 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
D Angle | -0.080594 | 1.963 | -0.74422 | -0.56911 | -0.034171 | -0.17875 | 0.09556 | -0.063389 | -0.045933 | 0.17063 | 0.094246 | -0.43177 | -0.11901 | 0.29689 | -0.59218 | -0.34895 | 0.57552 | 0.58903 | 0.63003 | 0.67559 | 0.51917
db2 | 0.45108 | -2.0638 | -0.5021 | -0.89368 | -0.31524 | 0.25506 | -0.10209 | -0.01422 | 0.63456 | -0.1746 | -0.66331 | 0.22869 | -0.21424 | 0.11344 | 0.46945 | 0.4095 | -0.13219 | 0.008293 | 0.43487 | 0.45781 | 0.14535
db3 | -0.68146 | 26.0355 | 0.69971 | -0.032177 | 0.11474 | -0.13667 | -0.22879 | -0.37676 | -0.2457 | 0.084885 | -0.32654 | -0.20588 | 0.049777 | 0.17868 | -0.14134 | 0.010772 | 0.089439 | 0.1512 | -0.031219 | 0.34588 | -5.2297
db4 | -0.0066973 | -1.3914 | 0.059769 | -0.023176 | -0.016458 | 0.019355 | -0.0702 | -0.042788 | 0.031045 | 0.0931 | 0.0040237 | -0.024453 | -0.031781 | -0.031396 | 0.049284 | -0.017277 | -0.050546 | 0.0095442 | -0.02398 | 0.065935 | -0.29276
db5 | -0.20204 | -5.7754 | -0.052981 | -0.15143 | -0.13718 | -0.1361 | 0.14546 | 0.080738 | 0.058618 | 0.071138 | 0.12533 | 0.01645 | -0.031663 | 0.053695 | 0.032017 | 0.012471 | -0.039846 | 0.070187 | -0.043116 | -0.073788 | -4.1581
db6 | 0.039071 | 0.50395 | 0.07861 | 0.017378 | -0.12484 | 0.050955 | 0.035941 | 0.043078 | -0.048858 | 0.073351 | 0.11197 | -0.077556 | 0.065984 | -0.10715 | -0.025461 | -0.055888 | 0.055192 | -0.096464 | -0.057489 | 0.061244 | -0.0050193
db7 | 0.9755 | 1.6647 | -0.10196 | 0.0029378 | 0.032081 | 0.0026463 | -0.02193 | 0.048089 | -0.0087657 | 0.01322 | -0.027246 | 0.00062602 | -0.0041254 | 0.035425 | 0.016859 | -0.051786 | 0.012051 | 0.049203 | -0.00081078 | 0.0034816 | 0.2716
db8 | 0.0071323 | 0.096574 | 0.05742 | 0.021968 | -0.15538 | 0.016104 | -0.044173 | -0.00092291 | -0.028364 | 0.034941 | 0.04448 | -0.093019 | 0.057814 | -0.078788 | 0.029716 | 0.017048 | 0.064887 | -0.03993 | 0.035752 | 0.060441 | -0.010899
db9 | 0.53529 | 0.44198 | 0.012057 | -0.015744 | 0.010195 | -0.0004776 | -0.0055068 | -0.0013103 | 0.0043951 | 0.0045636 | -0.0011701 | 0.0056763 | -0.0007802 | -0.001916 | -0.0051116 | -0.0022756 | -0.0036331 | 0.0057889 | -0.012856 | 0.0081058 | 0.45432
db10 | -0.009294 | 0.20918 | 0.07812 | 0.081212 | -0.2196 | 0.071591 | -0.029927 | 0.033597 | -0.064852 | 0.072675 | 0.13149 | -0.17288 | 0.089298 | -0.13764 | 0.047567 | 0.0051242 | 0.013575 | -0.10223 | 0.018397 | 0.08448 | 0.063669
db11 | 0.7758 | 0.60687 | 0.0034715 | -0.00054143 | -0.0022046 | -0.0027409 | -0.0019331 | -0.005877 | -0.0023513 | -0.00013713 | 0.0063654 | 0.0063791 | 0.0036056 | 0.0022155 | 0.0010727 | 0.0052896 | -0.0020418 | -0.0017961 | -0.004467 | -0.0043086 | 0.6318
db12 | -0.0011809 | 0.036897 | 0.0092758 | 0.0093889 | -0.031804 | 0.0087525 | -0.0064126 | 0.0031795 | -0.0049914 | 0.0069358 | 0.012276 | -0.022116 | 0.012703 | -0.018227 | 0.0085133 | 0.0043568 | 0.0035596 | -0.012455 | 0.0042938 | 0.012772 | 0.0049555
db13 | 0.084622 | 0.083063 | -0.0020729 | 0.0010044 | 0.0051869 | 0.0014686 | -0.00081395 | -0.00055656 | -0.0031653 | -0.0012146 | -0.0011537 | 0.0023971 | 0.00037619 | 0.0021195 | -0.00058695 | -0.00021276 | -0.00027762 | 0.0015914 | -0.0025581 | -0.0015316 | 0.052939
db14 | -0.0046997 | 0.0020249 | 0.0032524 | 0.0027384 | -0.0079497 | 0.0027844 | -0.0027151 | 0.0010826 | -0.0023548 | 0.0031961 | 0.0056602 | -0.0073651 | 0.0040831 | -0.0053792 | 0.0028486 | 0.00012868 | 0.00045782 | -0.0049593 | 0.00013009 | 0.0043607 | -0.0037957
db15 | 0.025862 | 0.002472 | 0.0049812 | 0.0017445 | -0.014243 | 0.0016588 | 0.00038316 | 0.002061 | -0.0021036 | 0.0023817 | 0.0061846 | -0.0090292 | 0.0032246 | -0.0064889 | 0.00064437 | -0.00039325 | 0.0022085 | -0.0012367 | 0.0055187 | 0.0025031 | 0.014158
da2 | 0.30025 | 9.5164 | -0.51775 | -0.038586 | -0.58258 | -0.18355 | -0.10712 | 0.54505 | 1.0528 | 0.75254 | 0.3177 | -0.39444 | -0.14149 | 0.31648 | 0.041562 | -0.1588 | -0.72634 | 0.18521 | -0.28389 | -0.076799 | -9.9728
da3 | -0.68659 | -3.0212 | 0.04364 | -0.17297 | -0.090732 | -0.1579 | -0.46652 | -0.66366 | 0.28936 | 0.1793 | 0.028462 | 0.45188 | 0.28666 | 0.24837 | 0.075304 | 0.037179 | 0.11021 | -0.37726 | -0.00015823 | 0.17885 | 0.14103
da4 | 0.14225 | 2.6065 | 0.063825 | -0.12998 | 0.068302 | -0.062573 | 0.2466 | 0.12382 | 0.16395 | 0.079373 | -0.17192 | -0.15834 | -0.025754 | -0.023173 | -0.11206 | -0.0647 | -0.014739 | 0.19525 | -0.086978 | -0.09089 | -0.15436
da5 | -0.13022 | 1.3367 | -0.0015841 | -0.056823 | 0.0032064 | -0.02972 | -0.018911 | -0.127 | 0.045541 | 0.10178 | 0.037006 | -0.0073939 | 0.041437 | 0.0074102 | 0.049537 | -0.008741 | -0.073703 | -0.042592 | 0.070219 | 0.010336 | -0.58608
da6 | 0.10005 | 0.033892 | -0.057569 | -0.010211 | 0.067562 | -0.0061048 | -0.0023978 | 0.00033066 | 0.11256 | 0.035753 | -0.0085066 | 0.053054 | -0.027131 | 0.076474 | 0.031195 | -0.016643 | -0.065141 | -0.053787 | -0.07905 | -0.050386 | -0.13684
da7 | 0.058008 | 1.7701 | 0.068289 | 0.013599 | -0.041505 | -0.033324 | -0.0063773 | -0.013303 | -0.005255 | 0.046786 | 0.023483 | -0.040763 | 0.00023948 | -0.010169 | 0.0071098 | 0.0064673 | 0.0126 | -0.012585 | -0.0099286 | -0.0053644 | 0.070396
da8 | -0.014563 | -0.046997 | -0.0042312 | 0.035832 | 0.099512 | -0.00017347 | 0.0031464 | -0.012755 | -0.028251 | -0.018323 | -0.055712 | 0.053501 | -0.023075 | 0.084746 | 0.032724 | 0.031302 | -0.047154 | -0.027853 | -0.075063 | -0.048173 | -0.074966
da9 | 0.051502 | -0.31141 | 0.0047181 | -0.0023362 | -0.050372 | -0.011465 | -0.0038686 | -0.011271 | 0.010886 | -0.005439 | -0.021368 | -0.00052201 | 0.0096066 | -0.0052063 | -0.0035361 | -0.012743 | 0.028007 | 0.029642 | 0.03735 | 0.0079189 | 0.018806
da10 | 0.0060104 | -0.069892 | -0.048228 | 0.046295 | 0.10981 | -0.016053 | -0.0146 | -0.029746 | -0.0068718 | -0.021275 | -0.070943 | 0.10252 | -0.052314 | 0.11836 | 0.047981 | 0.076326 | -0.027694 | -0.049942 | -0.1155 | -0.048126 | -0.10569
da11 | 0.063922 | 0.25846 | 0.0028407 | -0.0035852 | -0.0068294 | -0.003367 | 0.0046065 | 0.0017174 | -0.0040422 | 0.0096051 | 0.0015809 | -0.013675 | -0.0020434 | -0.0029236 | -0.0041952 | -0.0026718 | 0.0055485 | 0.0065348 | 0.0079105 | 0.0029885 | 0.06824
da12 | 0.0033505 | -0.023779 | -0.0078283 | 0.0044164 | 0.014477 | -0.0022545 | -0.0017988 | -0.0028243 | 0.0038174 | -0.0010444 | -0.0086867 | 0.013176 | -0.0081882 | 0.015928 | 0.0060021 | 0.0073656 | -0.0054883 | -0.0064064 | -0.014301 | -0.0063615 | -0.011596
da13 | 0.0072435 | -0.018959 | 0.00073893 | -0.00086546 | -0.002154 | -0.0022668 | -0.00074172 | -0.0034798 | 0.0009777 | -0.00021115 | -0.0011102 | 0.00082158 | 0.0013556 | 0.001159 | -0.0014833 | -0.0020202 | 0.0024257 | 0.0028491 | 0.0030787 | 0.00092622 | 0.0037508
da14 | -0.007366 | 0.011186 | -0.0023366 | 0.0015287 | 0.0053681 | -0.00027515 | 0.0011817 | 0.00018226 | -0.00014713 | -0.00053345 | -0.0047733 | 0.0044859 | -0.002831 | 0.0057753 | 0.0019924 | 0.0023378 | -0.0028454 | 0.00033959 | -0.0065916 | -0.0028581 | -0.004007
da15 | -0.000317 | 0.0096899 | 0.0014358 | 0.0021139 | 0.0002629 | 0.00022437 | -0.0039777 | -0.0027029 | -0.0025018 | 0.0017348 | 0.0032882 | 0.00091738 | 0.0017307 | 0.0019935 | 0.0014814 | 0.0011382 | 0.00041919 | -0.0071998 | -0.0024699 | 0.0021119 | -0.0050318
Check Computation for Aperture2
Magnetic Length
mbh dB/B | -3664.541
| mean 3:18 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
dB/B | 595.3216 | -3636.858 | 7.4151 | 1.483 | -0.49434 | -0.49434 | -0.49434 | -2.4717 | -2.4717 | 1.483 | -0.49434 | -0.49434 | -0.49434 | -0.49434 | 1.483 | 1.483 | -0.49434 | 1.483 | 1.483 | 7.4151 | -3692.224
| mean 2:19 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
D Angle | -0.13571 | 3.6582 | 0.49733 | 0.77247 | -0.11071 | -0.88987 | -0.091898 | -0.050869 | 0.73325 | 0.66824 | -0.61114 | -0.47788 | -0.028435 | -0.27227 | -0.40052 | -0.24436 | -0.14558 | 0.030824 | 0.12961 | 0.4918 | 0.44772
db2 | -0.73413 | -4.7971 | 0.93832 | 0.59063 | -0.83534 | 0.23908 | -0.0033491 | 0.57223 | -0.22027 | -0.058297 | -1.0599 | -0.7226 | -0.85014 | -0.32998 | 0.52674 | 0.7853 | 0.71169 | -0.78806 | -0.36383 | 0.8678 | -0.72161
db3 | -0.059674 | 28.936 | 0.14051 | -0.049976 | 0.18021 | -0.32976 | -0.48176 | -0.57951 | -0.11518 | 0.23236 | -0.69605 | -0.17904 | -0.21042 | 0.19814 | 0.2802 | 0.15325 | -0.15845 | 0.54965 | 0.34338 | 0.72245 | -4.6058
db4 | -0.080788 | -0.74965 | 0.0042796 | 0.070311 | -0.02729 | 0.023351 | -0.12582 | -0.15222 | 0.0061639 | 0.016613 | 0.035956 | -0.13546 | -0.063819 | 0.0097994 | -0.11834 | 0.091522 | 0.041563 | 0.15737 | 0.071543 | 0.094482 | 0.012733
db5 | -0.15984 | -5.7775 | -0.053769 | -0.036667 | 0.029562 | 0.11512 | 0.29964 | -0.047171 | -0.020742 | 0.033405 | 0.12145 | 0.012786 | -0.040096 | 0.030543 | 0.0064426 | 0.080264 | -0.11337 | 0.011996 | -0.11743 | -0.31197 | -3.4449
db6 | -0.075702 | -0.10598 | -0.057527 | 0.0088246 | 0.0079662 | -0.041305 | -0.040123 | 0.02911 | -0.1464 | -0.0098586 | 0.049916 | 0.011562 | 0.065172 | -0.045045 | 0.00023729 | 0.021624 | 0.069551 | -0.033878 | -0.020922 | 0.13109 | -0.010278
db7 | 1.0216 | 1.7078 | -0.10897 | 0.026848 | 0.0075377 | 0.044489 | -0.014713 | 0.031776 | -0.0012316 | 0.016016 | -0.02697 | 0.013303 | 0.023244 | 0.038153 | 0.04803 | 0.00025552 | -0.00037418 | -0.027601 | -0.043754 | -0.026039 | 0.17828
db8 | -0.0067732 | -0.071928 | 0.0067917 | -0.041586 | 0.023406 | -0.021406 | 0.0018373 | -0.022026 | -0.14113 | -0.032762 | 0.077669 | 0.016671 | 0.041416 | -0.030778 | -0.0022928 | 0.025843 | 0.025956 | 0.011362 | 0.011866 | 0.049163 | -0.097212
db9 | 0.54356 | 0.39137 | 0.0091578 | 0.0042259 | -0.002037 | -0.00041424 | -0.010397 | -0.0081743 | -0.0036076 | -0.012398 | 0.0032324 | 0.0046284 | 0.0020574 | 0.0080877 | -0.011806 | 0.0099508 | -0.0082692 | 0.0090086 | 0.0021669 | 0.0045886 | 0.4823
db10 | 0.001467 | 0.035514 | -0.028964 | -0.10755 | 0.022437 | -0.053707 | 0.028304 | -0.013209 | -0.20913 | -0.049498 | 0.098579 | 0.036735 | 0.058202 | -0.067641 | 0.014498 | 0.048896 | 0.067766 | 0.0093639 | 0.044859 | 0.10006 | -0.054805
db11 | 0.7718 | 0.58659 | 8.6444e-06 | 0.0012998 | -0.0020742 | 0.0022403 | 0.0018274 | -0.0049078 | -0.0033454 | -0.0014802 | 0.0026878 | -0.0028668 | -0.0032927 | -0.0015404 | -0.0029051 | 0.004994 | -0.00094186 | 0.0071494 | 0.0011508 | 0.0019956 | 0.65528
db12 | 0.0011853 | -0.0048772 | 0.00039196 | -0.012848 | 0.003157 | -0.0064903 | 0.0066538 | -0.0013076 | -0.029954 | -0.0080858 | 0.01236 | 0.0026755 | 0.0054786 | -0.0088325 | 0.0028593 | 0.0084342 | 0.010674 | -0.00098377 | 0.0029597 | 0.012858 | -0.014588
db13 | 0.086007 | 0.082833 | -0.0016113 | 0.0016582 | -0.0015349 | -0.0010968 | -0.0041759 | 0.0015105 | 0.00043038 | 0.00047267 | -0.0021971 | 0.00026854 | 0.0014017 | 0.00060696 | 0.00068573 | 0.00085384 | 0.0014698 | -0.0015541 | -0.00097401 | 0.0037858 | 0.047484
db14 | -0.003239 | 0.0024164 | -0.0023445 | -0.0045889 | 0.00012306 | -0.0041059 | 0.0015946 | -0.00012453 | -0.0074457 | -0.0014195 | 0.0045901 | 0.0016982 | 0.00263 | -0.0025753 | 0.0015768 | 0.0016652 | 0.0035958 | -0.0012879 | 0.0013928 | 0.0050255 | -0.0057776
db15 | 0.027606 | -0.00086662 | 0.002935 | -0.0041655 | 0.0048891 | 0.00057507 | 0.0063376 | 0.00019036 | -0.013905 | -0.0039412 | 0.0017431 | -0.00030866 | 0.00040669 | -0.0033425 | 0.00094023 | 0.0013932 | 0.00041594 | 0.0031158 | 0.0032499 | -0.00052909 | 0.014979
da2 | 1.4704 | 9.3715 | -0.1157 | -0.22974 | 0.0018752 | -0.10236 | -0.41626 | -0.052596 | 0.57024 | -0.16635 | -0.29351 | 0.048979 | -0.14571 | 0.057342 | -0.092087 | -0.16777 | 0.061139 | 0.34176 | 0.070588 | 0.63017 | -10.3166
da3 | -0.41776 | -0.69418 | 0.2418 | -0.058542 | -0.11163 | 0.25622 | -0.059473 | -0.077918 | 0.065476 | 0.18193 | -0.10465 | 0.51008 | 0.37573 | 0.50525 | 0.37345 | -0.45579 | -0.60392 | -0.21265 | -0.19766 | -0.62771 | -0.88559
da4 | 0.31281 | 1.5779 | -0.1469 | -0.22754 | -0.056698 | 0.32001 | 0.18396 | -0.26633 | -0.18981 | -0.088776 | 0.062592 | 0.11386 | 0.18282 | 0.10708 | 0.2967 | 0.030936 | -0.030705 | -0.04431 | -0.060632 | -0.18627 | -0.7786
da5 | -0.11982 | 1.6208 | 0.027486 | -0.015016 | -0.055672 | -0.010041 | 0.087114 | -0.047999 | -0.0042793 | 0.082246 | -0.050179 | 0.032519 | 0.10368 | 0.095252 | 0.16804 | -0.074351 | -0.065103 | -0.024794 | -0.07781 | -0.17109 | -0.56584
da6 | -0.010129 | -0.18941 | 0.034671 | -0.061599 | -0.077379 | 0.041079 | -0.015378 | -0.087275 | 0.019253 | 0.021746 | -0.010888 | 0.018373 | 0.017226 | 0.084519 | 0.099287 | -0.035942 | -0.01815 | -0.078156 | -0.0030267 | 0.051641 | 0.10453
da7 | 0.083848 | 1.8741 | -0.010721 | -0.034892 | -0.034224 | -0.0087183 | 0.083943 | 0.010581 | -0.019889 | 0.063155 | 0.040414 | 0.011331 | 0.026033 | 0.051905 | 0.023787 | -0.055193 | -0.058634 | -0.022852 | -0.019448 | -0.046578 | 0.17106
da8 | 0.020854 | -0.089879 | -0.04681 | -0.0022729 | -0.053058 | 0.0058967 | -0.020085 | -0.03848 | 0.10084 | 0.047107 | 0.0018365 | 0.019554 | -0.023659 | 0.033688 | 0.0047113 | -0.0039809 | -0.013703 | -0.023114 | 0.001189 | 0.010339 | -0.049248
da9 | 0.047154 | -0.22032 | 0.035779 | -0.020429 | -0.0056056 | -0.029396 | 0.0097375 | -0.00043938 | -0.023083 | 0.0082643 | -0.031226 | 0.022709 | -0.0046061 | 0.0068061 | -0.01372 | -0.01204 | 0.0093344 | 0.023929 | 0.02005 | 0.0039359 | -0.040857
da10 | 0.0066035 | -0.092078 | -0.026768 | 0.027905 | -0.086228 | -0.025148 | -0.062967 | -0.011887 | 0.16866 | 0.086464 | -0.0063646 | 0.0082251 | -0.020103 | 0.030747 | -0.017259 | -0.0093481 | 0.00839 | -0.046446 | -0.047955 | 0.030085 | -0.08841
da11 | 0.064894 | 0.26454 | 0.0026448 | 0.0021772 | -0.0059313 | -0.011306 | 0.012207 | -0.0013438 | -0.0017033 | 0.016224 | -0.0015171 | -0.0047132 | 0.0035827 | 0.00029612 | 0.0012932 | -0.0058608 | 0.00056838 | -0.0017253 | -0.0044537 | -0.00043882 | 0.092424
da12 | -0.0010344 | -0.017591 | -0.0023922 | -0.00049951 | -0.013323 | -0.0017693 | -0.0067041 | -0.0047272 | 0.01829 | 0.011575 | -0.0011792 | 0.0019209 | 0.00049543 | 0.0084638 | 0.00080014 | -0.0019656 | -0.00033363 | -0.0067353 | -0.0066144 | 0.0046978 | -0.0038417
da13 | 0.0076319 | -0.010616 | 0.0048935 | 0.00046163 | -0.00036396 | -0.003559 | 0.00018765 | -0.0013625 | 9.8467e-05 | 0.001818 | -0.0032892 | 0.0010019 | 0.0010708 | 0.0026906 | 0.00068444 | -0.002057 | -0.00022341 | 0.00084879 | -9.6318e-05 | -0.0028044 | -0.0012332
da14 | -0.00824 | 0.0068378 | -0.0030158 | 0.0010942 | -0.0045964 | 0.00016974 | -0.0024579 | -0.0012264 | 0.0063035 | 0.0038724 | 0.00054493 | 0.0017017 | -0.00033388 | 0.0021179 | -0.00023555 | -0.00018414 | -0.00025257 | -0.0023563 | -0.0021162 | 0.00097085 | -0.0039953
da15 | -0.00041557 | -0.0033195 | -0.0021753 | -0.0039041 | -0.0014179 | -0.0048504 | -0.00033978 | -0.0012057 | 0.0032348 | 0.0010226 | 0.0048733 | 0.0012312 | 0.0024526 | -0.0012621 | 0.00037808 | -9.1772e-06 | 0.00072055 | -0.00094914 | 0.00043536 | 0.0017653 | -0.015458
Field Colinearity | -0.33006
b1 |  | b3 |  | b5 |  | b7 |  | b9 |  | b11
inner | outer | inner | outer | inner | outer | inner | outer | inner | outer | inner | outer
59 | 40.2 | 21.8 | 16.2 | -4 | -0.77 | 1.5 | -0.22 | -0.56 | 0.031 | 0.083 | 0
Sensitivity to shims
New cross section bounds ? | yes

## Alstom Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | 14.45 | 0.01 | 3.5 | 7 | Field Colinearity |  | 0 | 0.32 | 3.5 | 7
dB/B Heads CS+NCS (units) | -3750 | 130 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 595.85 | 0.36 | 4 | 8 | 0 | 1.5 | 4 | 8 | -3848.5 | 151 | 3.5 | 7 | -3657.5 | 284 | 3.5 | 7 | 8.2 | 2 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | 0 | 0.5 | 4 | 8 | 5 | 2 | 3.5 | 7 | -0.4 | 1 | 3.5 | 7
b2 (units) | 0 | 0.64 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 0 | 5 | 3.5 | 7 | 0 | 2 | 3.5 | 7
b3 (units) | -4 | 2 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 39 | 4.6 | 3.5 | 7 | -11.57 | 2.5 | 3.5 | 7
b4 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | 0 | 0.9 | 3.5 | 7 | 0 | 0.38 | 3.5 | 7
b5 (units) | 0.04 | 0.5 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | -1.75 | 1.3 | 3.5 | 7 | -2.3 | 0.7 | 3.5 | 7
b6 (units) | 0 | 0.054 | 3.5 | 7 | 0 | 0.07 | 4 | 8 | 0 | 0.34 | 3.5 | 7 | 0 | 0.1 | 3.5 | 7
b7 (units) | 1.03 | 0.09 | 3.5 | 7 | 0 | 0.054 | 4 | 8 | 3.01 | 0.3 | 3.5 | 7 | 0.82 | 0.2 | 3.5 | 7
b8 (units) | 0 | 0.02 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.05 | 3.5 | 7
b9 (units) | 0.53 | 0.03 | 3.5 | 7 | 0 | 0.02 | 4 | 8 | 0.566 | 0.05 | 3.5 | 7 | 0.43 | 0.06 | 3.5 | 7
b10 (units) | 0 | 0.0045 | 3.5 | 7 | 0 | 0.04 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.075 | 3.5 | 7
b11 (units) | 0.74 | 0.011 | 3.5 | 7 | 0 | 0.0059 | 4 | 8 | 0.632 | 0.03 | 3.5 | 7 | 0.65 | 0.018 | 3.5 | 7
b12 (units) | 0 | 0.0018 | 3.5 | 7 | 0 | 0.0061 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.01 | 3.5 | 7
b13 (units) | 0.085 | 0.0083 | 3.5 | 7 | 0 | 0.0025 | 4 | 8 | 0.0924 | 0.01 | 3.5 | 7 | 0.059 | 0.009 | 3.5 | 7
b14 (units) | 0 | 0.006 | 3.5 | 7 | 0 | 0.0034 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | 0 | 0.0035 | 3.5 | 7
b15 (units) | 0.034 | 0.007 | 3.5 | 7 | 0 | 0.0027 | 4 | 8 | -0.0018 | 0.006 | 3.5 | 7 | 0.0129 | 0.0068 | 3.5 | 7
a2 (units) | 0 | 1 | 3.5 | 7 | 0 | 1.1 | 4 | 8 | 0 | 6 | 3.5 | 7 | 0 | 3 | 3.5 | 7
a3 (units) | -0.13 | 0.35 | 3.5 | 7 | 0 | 0.35 | 4 | 8 | -1.7 | 2.1 | 3.5 | 7 | 1.35 | 0.7 | 3.5 | 7
a4 (units) | 0 | 0.27 | 3.5 | 7 | 0 | 0.27 | 4 | 8 | -0.853 | 1.6 | 3.5 | 7 | 0 | 0.7 | 3.5 | 7
a5 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.12 | 4 | 8 | 2.4 | 0.83 | 3.5 | 7 | -0.18 | 0.25 | 3.5 | 7
a6 (units) | 0 | 0.12 | 3.5 | 7 | 0 | 0.068 | 4 | 8 | 0 | 0.4 | 3.5 | 7 | 0 | 0.17 | 3.5 | 7
a7 (units) | 0.005 | 0.03 | 3.5 | 7 | 0 | 0.056 | 4 | 8 | 1.9 | 0.25 | 3.5 | 7 | -0.0102 | 0.077 | 3.5 | 7
a8 (units) | -0.02 | 0.035 | 3.5 | 7 | 0 | 0.025 | 4 | 8 | 0 | 0.12 | 3.5 | 7 | 0 | 0.06 | 3.5 | 7
a9 (units) | 0.015 | 0.03 | 3.5 | 7 | 0 | 0.021 | 4 | 8 | -0.114 | 0.05 | 3.5 | 7 | 0 | 0.024 | 3.5 | 7
a10 (units) | 0.002 | 0.003 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.085 | 3.5 | 7 | 0 | 0.045 | 3.5 | 7
a11 (units) | 0.0127 | 0.03 | 3.5 | 7 | 0 | 0.009 | 4 | 8 | 0.2 | 0.025 | 3.5 | 7 | 0 | 0.026 | 3.5 | 7
a12 (units) | 0 | 0.005 | 3.5 | 7 | 0 | 0.0056 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
a13 (units) | 0 | 0.004 | 3.5 | 7 | 0 | 0.0022 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.0037 | 3.5 | 7
a14 (units) | -0.012 | 0.006 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | -0.003 | 0.004 | 3.5 | 7
a15 (units) | 0 | 0.005 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Magnetic Length (mm) | 14.415 | 14.485 | 14.38 | 14.52 | Field Colinearity |  | -1.12 | 1.12 | -2.24 | 2.24
dB/B Heads CS+NCS (units) | -4205 | -3295 | -4660 | -2840
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 594.41 | 597.29 | 592.97 | 598.73 | -6 | 6 | -12 | 12 | -4377 | -3320 | -4905.5 | -2791.5 | -4651.5 | -2663.5 | -5645.5 | -1669.5 | 1.2 | 15.2 | -5.8 | 22.2
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -2 | 2 | -4 | 4 | -2 | 12 | -9 | 19 | -3.9 | 3.1 | -7.4 | 6.6
b2 (units) | -2.24 | 2.24 | -4.48 | 4.48 | -2.4 | 2.4 | -4.8 | 4.8 | -17.5 | 17.5 | -35 | 35 | -7 | 7 | -14 | 14
b3 (units) | -11 | 3 | -18 | 10 | -2.4 | 2.4 | -4.8 | 4.8 | 22.9 | 55.1 | 6.8 | 71.2 | -20.32 | -2.82 | -29.07 | 5.93
b4 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.64 | 0.64 | -1.28 | 1.28 | -3.15 | 3.15 | -6.3 | 6.3 | -1.33 | 1.33 | -2.66 | 2.66
b5 (units) | -1.71 | 1.79 | -3.46 | 3.54 | -0.64 | 0.64 | -1.28 | 1.28 | -6.3 | 2.8 | -10.85 | 7.35 | -4.75 | 0.15 | -7.2 | 2.6
b6 (units) | -0.189 | 0.189 | -0.378 | 0.378 | -0.28 | 0.28 | -0.56 | 0.56 | -1.19 | 1.19 | -2.38 | 2.38 | -0.35 | 0.35 | -0.7 | 0.7
b7 (units) | 0.715 | 1.345 | 0.4 | 1.66 | -0.216 | 0.216 | -0.432 | 0.432 | 1.96 | 4.06 | 0.91 | 5.11 | 0.12 | 1.52 | -0.58 | 2.22
b8 (units) | -0.07 | 0.07 | -0.14 | 0.14 | -0.136 | 0.136 | -0.272 | 0.272 | -0.35 | 0.35 | -0.7 | 0.7 | -0.175 | 0.175 | -0.35 | 0.35
b9 (units) | 0.425 | 0.635 | 0.32 | 0.74 | -0.08 | 0.08 | -0.16 | 0.16 | 0.391 | 0.741 | 0.216 | 0.916 | 0.22 | 0.64 | 0.01 | 0.85
b10 (units) | -0.01575 | 0.01575 | -0.0315 | 0.0315 | -0.16 | 0.16 | -0.32 | 0.32 | -0.35 | 0.35 | -0.7 | 0.7 | -0.2625 | 0.2625 | -0.525 | 0.525
b11 (units) | 0.7015 | 0.7785 | 0.663 | 0.817 | -0.0236 | 0.0236 | -0.0472 | 0.0472 | 0.527 | 0.737 | 0.422 | 0.842 | 0.587 | 0.713 | 0.524 | 0.776
b12 (units) | -0.0063 | 0.0063 | -0.0126 | 0.0126 | -0.0244 | 0.0244 | -0.0488 | 0.0488 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.035 | 0.035 | -0.07 | 0.07
b13 (units) | 0.05595 | 0.11405 | 0.0269 | 0.1431 | -0.01 | 0.01 | -0.02 | 0.02 | 0.0574 | 0.1274 | 0.0224 | 0.1624 | 0.0275 | 0.0905 | -0.004 | 0.122
b14 (units) | -0.021 | 0.021 | -0.042 | 0.042 | -0.0136 | 0.0136 | -0.0272 | 0.0272 | -0.021 | 0.021 | -0.042 | 0.042 | -0.01225 | 0.01225 | -0.0245 | 0.0245
b15 (units) | 0.0095 | 0.0585 | -0.015 | 0.083 | -0.0108 | 0.0108 | -0.0216 | 0.0216 | -0.0228 | 0.0192 | -0.0438 | 0.0402 | -0.0109 | 0.0367 | -0.0347 | 0.0605
a2 (units) | -3.5 | 3.5 | -7 | 7 | -4.4 | 4.4 | -8.8 | 8.8 | -21 | 21 | -42 | 42 | -10.5 | 10.5 | -21 | 21
a3 (units) | -1.355 | 1.095 | -2.58 | 2.32 | -1.4 | 1.4 | -2.8 | 2.8 | -9.05 | 5.65 | -16.4 | 13 | -1.1 | 3.8 | -3.55 | 6.25
a4 (units) | -0.945 | 0.945 | -1.89 | 1.89 | -1.08 | 1.08 | -2.16 | 2.16 | -6.453 | 4.747 | -12.053 | 10.347 | -2.45 | 2.45 | -4.9 | 4.9
a5 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.48 | 0.48 | -0.96 | 0.96 | -0.505 | 5.305 | -3.41 | 8.21 | -1.055 | 0.695 | -1.93 | 1.57
a6 (units) | -0.42 | 0.42 | -0.84 | 0.84 | -0.272 | 0.272 | -0.544 | 0.544 | -1.4 | 1.4 | -2.8 | 2.8 | -0.595 | 0.595 | -1.19 | 1.19
a7 (units) | -0.1 | 0.11 | -0.205 | 0.215 | -0.224 | 0.224 | -0.448 | 0.448 | 1.025 | 2.775 | 0.15 | 3.65 | -0.2797 | 0.2593 | -0.5492 | 0.5288
a8 (units) | -0.1425 | 0.1025 | -0.265 | 0.225 | -0.1 | 0.1 | -0.2 | 0.2 | -0.42 | 0.42 | -0.84 | 0.84 | -0.21 | 0.21 | -0.42 | 0.42
a9 (units) | -0.09 | 0.12 | -0.195 | 0.225 | -0.084 | 0.084 | -0.168 | 0.168 | -0.289 | 0.061 | -0.464 | 0.236 | -0.084 | 0.084 | -0.168 | 0.168
a10 (units) | -0.0085 | 0.0125 | -0.019 | 0.023 | -0.136 | 0.136 | -0.272 | 0.272 | -0.2975 | 0.2975 | -0.595 | 0.595 | -0.1575 | 0.1575 | -0.315 | 0.315
a11 (units) | -0.0923 | 0.1177 | -0.1973 | 0.2227 | -0.036 | 0.036 | -0.072 | 0.072 | 0.1125 | 0.2875 | 0.025 | 0.375 | -0.091 | 0.091 | -0.182 | 0.182
a12 (units) | -0.0175 | 0.0175 | -0.035 | 0.035 | -0.0224 | 0.0224 | -0.0448 | 0.0448 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.021 | 0.021 | -0.042 | 0.042
a13 (units) | -0.014 | 0.014 | -0.028 | 0.028 | -0.0088 | 0.0088 | -0.0176 | 0.0176 | -0.035 | 0.035 | -0.07 | 0.07 | -0.01295 | 0.01295 | -0.0259 | 0.0259
a14 (units) | -0.033 | 0.009 | -0.054 | 0.03 | -0.012 | 0.012 | -0.024 | 0.024 | -0.021 | 0.021 | -0.042 | 0.042 | -0.017 | 0.011 | -0.031 | 0.025
a15 (units) | -0.0175 | 0.0175 | -0.035 | 0.035 | -0.012 | 0.012 | -0.024 | 0.024 | -0.035 | 0.035 | -0.07 | 0.07 | -0.021 | 0.021 | -0.042 | 0.042

## Ansaldo Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | 14.45 | 0.01 | 3.5 | 7 | Field Colinearity |  | 0 | 0.32 | 3.5 | 7
dB/B Heads CS+NCS (units) | -3750 | 130 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 595.85 | 0.36 | 4 | 8 | 0 | 1.5 | 4 | 8 | -3848.5 | 151 | 3.5 | 7 | -3657.5 | 284 | 3.5 | 7 | 8.2 | 2 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | 0 | 0.5 | 4 | 8 | 5 | 2 | 3.5 | 7 | -0.4 | 1 | 3.5 | 7
b2 (units) | 0 | 0.64 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 0 | 5 | 3.5 | 7 | 0 | 2 | 3.5 | 7
b3 (units) | -4 | 2 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 22.4 | 4.6 | 3.5 | 7 | -14.752 | 2.5 | 3.5 | 7
b4 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | 0 | 0.9 | 3.5 | 7 | 0 | 0.38 | 3.5 | 7
b5 (units) | 0.13 | 0.5 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | -4.87 | 1.3 | 3.5 | 7 | -2.3 | 0.7 | 3.5 | 7
b6 (units) | 0 | 0.054 | 3.5 | 7 | 0 | 0.07 | 4 | 8 | 0 | 0.34 | 3.5 | 7 | 0 | 0.1 | 3.5 | 7
b7 (units) | 0.84 | 0.09 | 3.5 | 7 | 0 | 0.054 | 4 | 8 | 2.02 | 0.3 | 3.5 | 7 | 0.1505 | 0.2 | 3.5 | 7
b8 (units) | 0 | 0.02 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.05 | 3.5 | 7
b9 (units) | 0.48 | 0.03 | 3.5 | 7 | 0 | 0.02 | 4 | 8 | 0.447 | 0.05 | 3.5 | 7 | 0.43 | 0.06 | 3.5 | 7
b10 (units) | 0 | 0.0045 | 3.5 | 7 | 0 | 0.04 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.075 | 3.5 | 7
b11 (units) | 0.74 | 0.011 | 3.5 | 7 | 0 | 0.0059 | 4 | 8 | 0.55 | 0.03 | 3.5 | 7 | 0.65 | 0.018 | 3.5 | 7
b12 (units) | 0 | 0.0018 | 3.5 | 7 | 0 | 0.0061 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.01 | 3.5 | 7
b13 (units) | 0.085 | 0.0083 | 3.5 | 7 | 0 | 0.0025 | 4 | 8 | 0.0807 | 0.01 | 3.5 | 7 | 0.059 | 0.009 | 3.5 | 7
b14 (units) | 0 | 0.006 | 3.5 | 7 | 0 | 0.0034 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | 0 | 0.0035 | 3.5 | 7
b15 (units) | 0.043 | 0.007 | 3.5 | 7 | 0 | 0.0027 | 4 | 8 | -0.00515 | 0.006 | 3.5 | 7 | 0.0244 | 0.0068 | 3.5 | 7
a2 (units) | 0 | 1 | 3.5 | 7 | 0 | 1.1 | 4 | 8 | 0 | 6 | 3.5 | 7 | 0 | 3 | 3.5 | 7
a3 (units) | -0.46 | 0.35 | 3.5 | 7 | 0 | 0.35 | 4 | 8 | -1.7 | 2.1 | 3.5 | 7 | -0.113 | 0.7 | 3.5 | 7
a4 (units) | 0 | 0.27 | 3.5 | 7 | 0 | 0.27 | 4 | 8 | 2.02 | 1.6 | 3.5 | 7 | 0 | 0.7 | 3.5 | 7
a5 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.12 | 4 | 8 | 2.4 | 0.83 | 3.5 | 7 | -0.18 | 0.25 | 3.5 | 7
a6 (units) | 0 | 0.12 | 3.5 | 7 | 0 | 0.068 | 4 | 8 | 0 | 0.4 | 3.5 | 7 | 0 | 0.17 | 3.5 | 7
a7 (units) | 0.044 | 0.03 | 3.5 | 7 | 0 | 0.056 | 4 | 8 | 1.9 | 0.25 | 3.5 | 7 | 0.108 | 0.077 | 3.5 | 7
a8 (units) | -0.02 | 0.035 | 3.5 | 7 | 0 | 0.025 | 4 | 8 | 0 | 0.12 | 3.5 | 7 | 0 | 0.06 | 3.5 | 7
a9 (units) | 0.035 | 0.03 | 3.5 | 7 | 0 | 0.021 | 4 | 8 | -0.243 | 0.05 | 3.5 | 7 | 0 | 0.024 | 3.5 | 7
a10 (units) | 0.002 | 0.003 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.085 | 3.5 | 7 | 0 | 0.045 | 3.5 | 7
a11 (units) | 0.037 | 0.03 | 3.5 | 7 | 0 | 0.009 | 4 | 8 | 0.2 | 0.025 | 3.5 | 7 | 0 | 0.026 | 3.5 | 7
a12 (units) | 0 | 0.005 | 3.5 | 7 | 0 | 0.0056 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
a13 (units) | 0 | 0.004 | 3.5 | 7 | 0 | 0.0022 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.0037 | 3.5 | 7
a14 (units) | -0.012 | 0.006 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | -0.003 | 0.004 | 3.5 | 7
a15 (units) | 0.0058 | 0.005 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Magnetic Length (mm) | 14.415 | 14.485 | 14.38 | 14.52 | Field Colinearity |  | -1.12 | 1.12 | -2.24 | 2.24
dB/B Heads CS+NCS (units) | -4205 | -3295 | -4660 | -2840
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 594.41 | 597.29 | 592.97 | 598.73 | -6 | 6 | -12 | 12 | -4377 | -3320 | -4905.5 | -2791.5 | -4651.5 | -2663.5 | -5645.5 | -1669.5 | 1.2 | 15.2 | -5.8 | 22.2
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -2 | 2 | -4 | 4 | -2 | 12 | -9 | 19 | -3.9 | 3.1 | -7.4 | 6.6
b2 (units) | -2.24 | 2.24 | -4.48 | 4.48 | -2.4 | 2.4 | -4.8 | 4.8 | -17.5 | 17.5 | -35 | 35 | -7 | 7 | -14 | 14
b3 (units) | -11 | 3 | -18 | 10 | -2.4 | 2.4 | -4.8 | 4.8 | 6.3 | 38.5 | -9.8 | 54.6 | -23.502 | -6.002 | -32.252 | 2.748
b4 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.64 | 0.64 | -1.28 | 1.28 | -3.15 | 3.15 | -6.3 | 6.3 | -1.33 | 1.33 | -2.66 | 2.66
b5 (units) | -1.62 | 1.88 | -3.37 | 3.63 | -0.64 | 0.64 | -1.28 | 1.28 | -9.42 | -0.32 | -13.97 | 4.23 | -4.75 | 0.15 | -7.2 | 2.6
b6 (units) | -0.189 | 0.189 | -0.378 | 0.378 | -0.28 | 0.28 | -0.56 | 0.56 | -1.19 | 1.19 | -2.38 | 2.38 | -0.35 | 0.35 | -0.7 | 0.7
b7 (units) | 0.525 | 1.155 | 0.21 | 1.47 | -0.216 | 0.216 | -0.432 | 0.432 | 0.97 | 3.07 | -0.08 | 4.12 | -0.5495 | 0.8505 | -1.2495 | 1.5505
b8 (units) | -0.07 | 0.07 | -0.14 | 0.14 | -0.136 | 0.136 | -0.272 | 0.272 | -0.35 | 0.35 | -0.7 | 0.7 | -0.175 | 0.175 | -0.35 | 0.35
b9 (units) | 0.375 | 0.585 | 0.27 | 0.69 | -0.08 | 0.08 | -0.16 | 0.16 | 0.272 | 0.622 | 0.097 | 0.797 | 0.22 | 0.64 | 0.01 | 0.85
b10 (units) | -0.01575 | 0.01575 | -0.0315 | 0.0315 | -0.16 | 0.16 | -0.32 | 0.32 | -0.35 | 0.35 | -0.7 | 0.7 | -0.2625 | 0.2625 | -0.525 | 0.525
b11 (units) | 0.7015 | 0.7785 | 0.663 | 0.817 | -0.0236 | 0.0236 | -0.0472 | 0.0472 | 0.445 | 0.655 | 0.34 | 0.76 | 0.587 | 0.713 | 0.524 | 0.776
b12 (units) | -0.0063 | 0.0063 | -0.0126 | 0.0126 | -0.0244 | 0.0244 | -0.0488 | 0.0488 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.035 | 0.035 | -0.07 | 0.07
b13 (units) | 0.05595 | 0.11405 | 0.0269 | 0.1431 | -0.01 | 0.01 | -0.02 | 0.02 | 0.0457 | 0.1157 | 0.0107 | 0.1507 | 0.0275 | 0.0905 | -0.004 | 0.122
b14 (units) | -0.021 | 0.021 | -0.042 | 0.042 | -0.0136 | 0.0136 | -0.0272 | 0.0272 | -0.021 | 0.021 | -0.042 | 0.042 | -0.01225 | 0.01225 | -0.0245 | 0.0245
b15 (units) | 0.0185 | 0.0675 | -0.006 | 0.092 | -0.0108 | 0.0108 | -0.0216 | 0.0216 | -0.02615 | 0.01585 | -0.04715 | 0.03685 | 0.0006 | 0.0482 | -0.0232 | 0.072
a2 (units) | -3.5 | 3.5 | -7 | 7 | -4.4 | 4.4 | -8.8 | 8.8 | -21 | 21 | -42 | 42 | -10.5 | 10.5 | -21 | 21
a3 (units) | -1.685 | 0.765 | -2.91 | 1.99 | -1.4 | 1.4 | -2.8 | 2.8 | -9.05 | 5.65 | -16.4 | 13 | -2.563 | 2.337 | -5.013 | 4.787
a4 (units) | -0.945 | 0.945 | -1.89 | 1.89 | -1.08 | 1.08 | -2.16 | 2.16 | -3.58 | 7.62 | -9.18 | 13.22 | -2.45 | 2.45 | -4.9 | 4.9
a5 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.48 | 0.48 | -0.96 | 0.96 | -0.505 | 5.305 | -3.41 | 8.21 | -1.055 | 0.695 | -1.93 | 1.57
a6 (units) | -0.42 | 0.42 | -0.84 | 0.84 | -0.272 | 0.272 | -0.544 | 0.544 | -1.4 | 1.4 | -2.8 | 2.8 | -0.595 | 0.595 | -1.19 | 1.19
a7 (units) | -0.061 | 0.149 | -0.166 | 0.254 | -0.224 | 0.224 | -0.448 | 0.448 | 1.025 | 2.775 | 0.15 | 3.65 | -0.1615 | 0.3775 | -0.431 | 0.647
a8 (units) | -0.1425 | 0.1025 | -0.265 | 0.225 | -0.1 | 0.1 | -0.2 | 0.2 | -0.42 | 0.42 | -0.84 | 0.84 | -0.21 | 0.21 | -0.42 | 0.42
a9 (units) | -0.07 | 0.14 | -0.175 | 0.245 | -0.084 | 0.084 | -0.168 | 0.168 | -0.418 | -0.068 | -0.593 | 0.107 | -0.084 | 0.084 | -0.168 | 0.168
a10 (units) | -0.0085 | 0.0125 | -0.019 | 0.023 | -0.136 | 0.136 | -0.272 | 0.272 | -0.2975 | 0.2975 | -0.595 | 0.595 | -0.1575 | 0.1575 | -0.315 | 0.315
a11 (units) | -0.068 | 0.142 | -0.173 | 0.247 | -0.036 | 0.036 | -0.072 | 0.072 | 0.1125 | 0.2875 | 0.025 | 0.375 | -0.091 | 0.091 | -0.182 | 0.182
a12 (units) | -0.0175 | 0.0175 | -0.035 | 0.035 | -0.0224 | 0.0224 | -0.0448 | 0.0448 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.021 | 0.021 | -0.042 | 0.042
a13 (units) | -0.014 | 0.014 | -0.028 | 0.028 | -0.0088 | 0.0088 | -0.0176 | 0.0176 | -0.035 | 0.035 | -0.07 | 0.07 | -0.01295 | 0.01295 | -0.0259 | 0.0259
a14 (units) | -0.033 | 0.009 | -0.054 | 0.03 | -0.012 | 0.012 | -0.024 | 0.024 | -0.021 | 0.021 | -0.042 | 0.042 | -0.017 | 0.011 | -0.031 | 0.025
a15 (units) | -0.0117 | 0.0233 | -0.0292 | 0.0408 | -0.012 | 0.012 | -0.024 | 0.024 | -0.035 | 0.035 | -0.07 | 0.07 | -0.021 | 0.021 | -0.042 | 0.042

## Noell Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | 14.45 | 0.01 | 3.5 | 7 | Field Colinearity |  | 0 | 0.32 | 3.5 | 7
dB/B Heads CS+NCS (units) | -3750 | 130 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 595.85 | 0.36 | 4 | 8 | 0 | 1.5 | 4 | 8 | -3848.5 | 151 | 3.5 | 7 | -3657.5 | 284 | 3.5 | 7 | 8.2 | 2 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | 0 | 0.5 | 4 | 8 | 5 | 2 | 3.5 | 7 | -0.4 | 1 | 3.5 | 7
b2 (units) | 0 | 0.64 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 0 | 5 | 3.5 | 7 | 0 | 2 | 3.5 | 7
b3 (units) | -4 | 2 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 32.9 | 4.6 | 3.5 | 7 | -5.67 | 2.5 | 3.5 | 7
b4 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | 0 | 0.9 | 3.5 | 7 | 0 | 0.38 | 3.5 | 7
b5 (units) | -0.56 | 0.5 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | -3.74 | 1.3 | 3.5 | 7 | -2.3 | 0.7 | 3.5 | 7
b6 (units) | 0 | 0.054 | 3.5 | 7 | 0 | 0.07 | 4 | 8 | 0 | 0.34 | 3.5 | 7 | 0 | 0.1 | 3.5 | 7
b7 (units) | 1 | 0.09 | 3.5 | 7 | 0 | 0.054 | 4 | 8 | 2.94 | 0.3 | 3.5 | 7 | 0.589 | 0.2 | 3.5 | 7
b8 (units) | 0 | 0.02 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.05 | 3.5 | 7
b9 (units) | 0.46 | 0.03 | 3.5 | 7 | 0 | 0.02 | 4 | 8 | 0.582 | 0.05 | 3.5 | 7 | 0.43 | 0.06 | 3.5 | 7
b10 (units) | 0 | 0.0045 | 3.5 | 7 | 0 | 0.04 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.075 | 3.5 | 7
b11 (units) | 0.74 | 0.011 | 3.5 | 7 | 0 | 0.0059 | 4 | 8 | 0.603 | 0.03 | 3.5 | 7 | 0.65 | 0.018 | 3.5 | 7
b12 (units) | 0 | 0.0018 | 3.5 | 7 | 0 | 0.0061 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.01 | 3.5 | 7
b13 (units) | 0.085 | 0.0083 | 3.5 | 7 | 0 | 0.0025 | 4 | 8 | 0.107 | 0.01 | 3.5 | 7 | 0.059 | 0.009 | 3.5 | 7
b14 (units) | 0 | 0.006 | 3.5 | 7 | 0 | 0.0034 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | 0 | 0.0035 | 3.5 | 7
b15 (units) | 0.032 | 0.007 | 3.5 | 7 | 0 | 0.0027 | 4 | 8 | -0.008 | 0.006 | 3.5 | 7 | 0.0135 | 0.0068 | 3.5 | 7
a2 (units) | 0 | 1 | 3.5 | 7 | 0 | 1.1 | 4 | 8 | 0 | 6 | 3.5 | 7 | 0 | 3 | 3.5 | 7
a3 (units) | 0.12 | 0.35 | 3.5 | 7 | 0 | 0.35 | 4 | 8 | -1.7 | 2.1 | 3.5 | 7 | 0.988 | 0.7 | 3.5 | 7
a4 (units) | 0 | 0.27 | 3.5 | 7 | 0 | 0.27 | 4 | 8 | 0.76 | 1.6 | 3.5 | 7 | 0 | 0.7 | 3.5 | 7
a5 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.12 | 4 | 8 | 2.4 | 0.83 | 3.5 | 7 | -0.18 | 0.25 | 3.5 | 7
a6 (units) | 0 | 0.12 | 3.5 | 7 | 0 | 0.068 | 4 | 8 | 0 | 0.4 | 3.5 | 7 | 0 | 0.17 | 3.5 | 7
a7 (units) | 0.002 | 0.03 | 3.5 | 7 | 0 | 0.056 | 4 | 8 | 1.9 | 0.25 | 3.5 | 7 | -0.00256 | 0.077 | 3.5 | 7
a8 (units) | -0.02 | 0.035 | 3.5 | 7 | 0 | 0.025 | 4 | 8 | 0 | 0.12 | 3.5 | 7 | 0 | 0.06 | 3.5 | 7
a9 (units) | 0.0072 | 0.03 | 3.5 | 7 | 0 | 0.021 | 4 | 8 | -0.231 | 0.05 | 3.5 | 7 | 0 | 0.024 | 3.5 | 7
a10 (units) | 0.002 | 0.003 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.085 | 3.5 | 7 | 0 | 0.045 | 3.5 | 7
a11 (units) | 0.002 | 0.03 | 3.5 | 7 | 0 | 0.009 | 4 | 8 | 0.2 | 0.025 | 3.5 | 7 | 0 | 0.026 | 3.5 | 7
a12 (units) | 0 | 0.005 | 3.5 | 7 | 0 | 0.0056 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
a13 (units) | 0 | 0.004 | 3.5 | 7 | 0 | 0.0022 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.0037 | 3.5 | 7
a14 (units) | -0.012 | 0.006 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | -0.003 | 0.004 | 3.5 | 7
a15 (units) | 0 | 0.005 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Magnetic Length (mm) | 14.415 | 14.485 | 14.38 | 14.52 | Field Colinearity |  | -1.12 | 1.12 | -2.24 | 2.24
dB/B Heads CS+NCS (units) | -4205 | -3295 | -4660 | -2840
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 594.41 | 597.29 | 592.97 | 598.73 | -6 | 6 | -12 | 12 | -4377 | -3320 | -4905.5 | -2791.5 | -4651.5 | -2663.5 | -5645.5 | -1669.5 | 1.2 | 15.2 | -5.8 | 22.2
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -2 | 2 | -4 | 4 | -2 | 12 | -9 | 19 | -3.9 | 3.1 | -7.4 | 6.6
b2 (units) | -2.24 | 2.24 | -4.48 | 4.48 | -2.4 | 2.4 | -4.8 | 4.8 | -17.5 | 17.5 | -35 | 35 | -7 | 7 | -14 | 14
b3 (units) | -11 | 3 | -18 | 10 | -2.4 | 2.4 | -4.8 | 4.8 | 16.8 | 49 | 0.7 | 65.1 | -14.42 | 3.08 | -23.17 | 11.83
b4 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.64 | 0.64 | -1.28 | 1.28 | -3.15 | 3.15 | -6.3 | 6.3 | -1.33 | 1.33 | -2.66 | 2.66
b5 (units) | -2.31 | 1.19 | -4.06 | 2.94 | -0.64 | 0.64 | -1.28 | 1.28 | -8.29 | 0.81 | -12.84 | 5.36 | -4.75 | 0.15 | -7.2 | 2.6
b6 (units) | -0.189 | 0.189 | -0.378 | 0.378 | -0.28 | 0.28 | -0.56 | 0.56 | -1.19 | 1.19 | -2.38 | 2.38 | -0.35 | 0.35 | -0.7 | 0.7
b7 (units) | 0.685 | 1.315 | 0.37 | 1.63 | -0.216 | 0.216 | -0.432 | 0.432 | 1.89 | 3.99 | 0.84 | 5.04 | -0.111 | 1.289 | -0.811 | 1.989
b8 (units) | -0.07 | 0.07 | -0.14 | 0.14 | -0.136 | 0.136 | -0.272 | 0.272 | -0.35 | 0.35 | -0.7 | 0.7 | -0.175 | 0.175 | -0.35 | 0.35
b9 (units) | 0.355 | 0.565 | 0.25 | 0.67 | -0.08 | 0.08 | -0.16 | 0.16 | 0.407 | 0.757 | 0.232 | 0.932 | 0.22 | 0.64 | 0.01 | 0.85
b10 (units) | -0.01575 | 0.01575 | -0.0315 | 0.0315 | -0.16 | 0.16 | -0.32 | 0.32 | -0.35 | 0.35 | -0.7 | 0.7 | -0.2625 | 0.2625 | -0.525 | 0.525
b11 (units) | 0.7015 | 0.7785 | 0.663 | 0.817 | -0.0236 | 0.0236 | -0.0472 | 0.0472 | 0.498 | 0.708 | 0.393 | 0.813 | 0.587 | 0.713 | 0.524 | 0.776
b12 (units) | -0.0063 | 0.0063 | -0.0126 | 0.0126 | -0.0244 | 0.0244 | -0.0488 | 0.0488 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.035 | 0.035 | -0.07 | 0.07
b13 (units) | 0.05595 | 0.11405 | 0.0269 | 0.1431 | -0.01 | 0.01 | -0.02 | 0.02 | 0.072 | 0.142 | 0.037 | 0.177 | 0.0275 | 0.0905 | -0.004 | 0.122
b14 (units) | -0.021 | 0.021 | -0.042 | 0.042 | -0.0136 | 0.0136 | -0.0272 | 0.0272 | -0.021 | 0.021 | -0.042 | 0.042 | -0.01225 | 0.01225 | -0.0245 | 0.0245
b15 (units) | 0.0075 | 0.0565 | -0.017 | 0.081 | -0.0108 | 0.0108 | -0.0216 | 0.0216 | -0.029 | 0.013 | -0.05 | 0.034 | -0.0103 | 0.0373 | -0.0341 | 0.0611
a2 (units) | -3.5 | 3.5 | -7 | 7 | -4.4 | 4.4 | -8.8 | 8.8 | -21 | 21 | -42 | 42 | -10.5 | 10.5 | -21 | 21
a3 (units) | -1.105 | 1.345 | -2.33 | 2.57 | -1.4 | 1.4 | -2.8 | 2.8 | -9.05 | 5.65 | -16.4 | 13 | -1.462 | 3.438 | -3.912 | 5.888
a4 (units) | -0.945 | 0.945 | -1.89 | 1.89 | -1.08 | 1.08 | -2.16 | 2.16 | -4.84 | 6.36 | -10.44 | 11.96 | -2.45 | 2.45 | -4.9 | 4.9
a5 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.48 | 0.48 | -0.96 | 0.96 | -0.505 | 5.305 | -3.41 | 8.21 | -1.055 | 0.695 | -1.93 | 1.57
a6 (units) | -0.42 | 0.42 | -0.84 | 0.84 | -0.272 | 0.272 | -0.544 | 0.544 | -1.4 | 1.4 | -2.8 | 2.8 | -0.595 | 0.595 | -1.19 | 1.19
a7 (units) | -0.103 | 0.107 | -0.208 | 0.212 | -0.224 | 0.224 | -0.448 | 0.448 | 1.025 | 2.775 | 0.15 | 3.65 | -0.27206 | 0.26694 | -0.54156 | 0.53644
a8 (units) | -0.1425 | 0.1025 | -0.265 | 0.225 | -0.1 | 0.1 | -0.2 | 0.2 | -0.42 | 0.42 | -0.84 | 0.84 | -0.21 | 0.21 | -0.42 | 0.42
a9 (units) | -0.0978 | 0.1122 | -0.2028 | 0.2172 | -0.084 | 0.084 | -0.168 | 0.168 | -0.406 | -0.056 | -0.581 | 0.119 | -0.084 | 0.084 | -0.168 | 0.168
a10 (units) | -0.0085 | 0.0125 | -0.019 | 0.023 | -0.136 | 0.136 | -0.272 | 0.272 | -0.2975 | 0.2975 | -0.595 | 0.595 | -0.1575 | 0.1575 | -0.315 | 0.315
a11 (units) | -0.103 | 0.107 | -0.208 | 0.212 | -0.036 | 0.036 | -0.072 | 0.072 | 0.1125 | 0.2875 | 0.025 | 0.375 | -0.091 | 0.091 | -0.182 | 0.182
a12 (units) | -0.0175 | 0.0175 | -0.035 | 0.035 | -0.0224 | 0.0224 | -0.0448 | 0.0448 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.021 | 0.021 | -0.042 | 0.042
a13 (units) | -0.014 | 0.014 | -0.028 | 0.028 | -0.0088 | 0.0088 | -0.0176 | 0.0176 | -0.035 | 0.035 | -0.07 | 0.07 | -0.01295 | 0.01295 | -0.0259 | 0.0259
a14 (units) | -0.033 | 0.009 | -0.054 | 0.03 | -0.012 | 0.012 | -0.024 | 0.024 | -0.021 | 0.021 | -0.042 | 0.042 | -0.017 | 0.011 | -0.031 | 0.025
a15 (units) | -0.0175 | 0.0175 | -0.035 | 0.035 | -0.012 | 0.012 | -0.024 | 0.024 | -0.035 | 0.035 | -0.07 | 0.07 | -0.021 | 0.021 | -0.042 | 0.042

## MTF
Straight part | [21 empty] | Integral
Aperture 1 | [10 empty] | Aperture 2 | [10 empty] | Aperture 1 | [9 empty] | Aperture 2
b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Coil pos. | TF | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Coil pos. | TF | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Mag len | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Mag len
0.45108 | -2.8615 | -0.0066973 | 0.19796 | 0.8255 | -0.80594 | 0.30025 | -0.68659 | 0.14225 | 13.299 | 0.59507 | -0.73413 | -2.2397 | -0.080788 | 0.24016 | 0.87157 | -1.3571 | 1.4704 | -0.41776 | 0.31281 | 16.155 | 0.59502 | 0.35813 | -1.9861 | -0.061722 | -0.1417 | 0.83502 | 0 | 0.26762 | -0.73644 | 0.21378 | 14447.859 | -0.86781 | -1.2876 | -0.099805 | -0.079228 | 0.87648 | -1.0386e-13 | 1.3455 | -0.44223 | 0.31885 | 14448.382
Shims
Aperture 1 | [7 empty] | Aperture 2
Upper |  |  |  | Lower |  |  |  | Upper |  |  |  | Lower
Inner layer |  | Outer layer |  | Inner layer |  | Outer layer |  | Inner layer |  | Outer layer |  | Inner layer |  | Outer layer
Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central
0.1 | 0.1 | 0.8 | 0.8 | 0.1 | 0.1 | 0.8 | 0.8 | 0.1 | 0.1 | 0.8 | 0.8 | 0.1 | 0.1 | 0.8 | 0.8