## Alarm sheet
File name | HCMBB_A001-02000013_cc.xls
Component ID | HCMBB_A001 |  |  |  | Serial Number | 2000013
Date of test Ap 1 | 2002-06-26 |  |  |  | Date of test Ap 2 | 2002-06-26
| Aperture 1 |  |  |  |  | Aperture 2
Magnetic length | status ok |  |  |  | Magnetic length | status ok
| Average straight | Variation straight | Heads CS | Heads NCS |  | Average straight | Variation straight | Heads CS | Heads NCS
| positions 2 to 19 | positions 2 to 19 | position 1 | position 20 |  | positions 2 to 19 | positions 2 to 19 | position 1 | position 20
Main field | status ok | status ok | status ok | status ok | Main field | status ok | status ok | status ok | status ok
Angle |  | status ok | status ok | status ok | Angle |  | status ok | status ok | status ok
b2 | status ok | status ok | status ok | status ok | b2 | status ok | status ok | status ok | status ok
b3 | status ok | status ok | status ok | status ok | b3 | status ok | status ok | status ok | status ok
b4 | status ok | status ok | status ok | status ok | b4 | status ok | status ok | status ok | status ok
b5 | status ok | status ok | status ok | status ok | b5 | status ok | status ok | status ok | status ok
b6 | status ok | status ok | status ok | status ok | b6 | status ok | status ok | status ok | status ok
b7 | status ok | status ok | status ok | status ok | b7 | status ok | status ok | status ok | status ok
b8 | status ok | status ok | status ok | status ok | b8 | status ok | status ok | status ok | status ok
b9 | status ok | status ok | status ok | status ok | b9 | status ok | status ok | status ok | status ok
b10 | status ok | status ok | status ok | status ok | b10 | status ok | status ok | status ok | status ok
b11 | status ok | status ok | status ok | status ok | b11 | status ok | status ok | status ok | status ok
b12 | status ok | status ok | status ok | status ok | b12 | status ok | status ok | status ok | status ok
b13 | status ok | status ok | status ok | status ok | b13 | status ok | status ok | status ok | status ok
b14 | status ok | status ok | status ok | status ok | b14 | status ok | status ok | status ok | yellow alarm
b15 | status ok | status ok | yellow alarm | status ok | b15 | status ok | status ok | yellow alarm | status ok
a2 | status ok | status ok | status ok | status ok | a2 | status ok | status ok | status ok | status ok
a3 | status ok | status ok | status ok | status ok | a3 | status ok | status ok | status ok | status ok
a4 | status ok | status ok | status ok | status ok | a4 | status ok | status ok | status ok | status ok
a5 | status ok | status ok | status ok | status ok | a5 | status ok | status ok | status ok | status ok
a6 | status ok | status ok | status ok | status ok | a6 | status ok | status ok | status ok | status ok
a7 | status ok | status ok | status ok | status ok | a7 | status ok | status ok | status ok | status ok
a8 | status ok | status ok | status ok | status ok | a8 | status ok | status ok | status ok | status ok
a9 | status ok | status ok | status ok | status ok | a9 | status ok | status ok | status ok | status ok
a10 | status ok | status ok | status ok | status ok | a10 | status ok | status ok | status ok | status ok
a11 | status ok | status ok | status ok | status ok | a11 | status ok | status ok | status ok | status ok
a12 | status ok | status ok | status ok | status ok | a12 | status ok | status ok | status ok | status ok
a13 | status ok | status ok | yellow alarm | status ok | a13 | status ok | status ok | yellow alarm | status ok
a14 | status ok | status ok | status ok | status ok | a14 | status ok | status ok | status ok | status ok
a15 | yellow alarm | status ok | status ok | status ok | a15 | yellow alarm | status ok | status ok | yellow alarm
Coil Positioning | status ok |  |  |  | Coil Positioning | status ok
Field Colinearity | status ok

## Assembly Data
Prog. # cc | Magnet name |  |  | Coil cross-section |  |  | End Spacers
33 | HCMBB_A001 | 2000013 |  | 1 |  |  | 1 | Not filed
COMMENTS:
| HOLDING POINT | Ok | Data at CERN |  | 2002-05-22 | Analysis by |  | E. Todesco | Answer to MMS-LD |  |  | 2002-05-23 | Answer to manufacturer |  |  | 2002-05-23
| Shims-summary | [6 empty] | Shims
| Inner average |  |  | Outer average |  |  |  | Aperture 1 |  |  |  |  | Aperture 2
| Used | Nominal | Differ. | Used | Nominal | Differ. |  | Left |  | Right |  |  | Left |  | Right
| 0.2 | 0.2 | 0 | 0.8 | 0.8 | 0 |  | Outer | Inner | Inner | Outer |  | Outer | Inner | Inner | Outer
[8 empty] | 0.8 | 0.2 | 0.2 | 0.8 | Up | 0.8 | 0.2 | 0.2 | 0.8
| Pre-stress (estimate, Mpa) | [6 empty] | 0.8 | 0.2 | 0.2 | 0.8 | Down | 0.8 | 0.2 | 0.2 | 0.8
| Inner average |  |  | Outer average |  |  |  | Seen from conection side
| Used | Nominal | Differ. | Used | Nominal | Differ.
| 75 | 75 | 0 | 75 | 75 | 0
| Peak pre-stress during collaring |  |  |  |  | 110
| Collared coil template - version 08.04.02 | [8 empty] | Version : |  | CERN
1 | Remplir les champs jaunes dans "Assembly Data" et dans "Original Data" | [12 empty] | French
2 | CTRL-g pour charger les fichers des mesures .txt
3 | Sauver le fichier avec un nom different
4 | ALT-F8 et appuyer sur "collared coil analysis" pour analyser les mesures - resultats dans Alarm sheet
1 | Fuellen Sie die gelben Felder in "Assembly Data" und in "Original Data" | [12 empty] | German
2 | CTRL-g um die .txt Messdatensaetze zu laden
3 | Sichern Sie diesen Datensatz mit einem anderen Namen
4 | ALT-F8 und auf "collared coil analysis" druecken, um die Daten zu analysieren => Ergebnis in "Alarm sheet"
1 | Riempire I campi gialli in "Assembly Data" e in "Original Data" | [12 empty] | Italian
2 | CTRL-g per caricare I file .txt delle misure
3 | Salvare il file con un nome diverso
4 | ALT-F8 e premere su "collared coil analysis" per analiszzare le misure - risultati in "Alarm sheet"
1 | Fill in the yellow fields in "Assembly Data" and "Original Data" | [12 empty] | English
2 | CTRL-g to load the .txt files of the measurements
3 | Save the file with a different name
4 | ALT-F8 and press on "collared coil analysis" to analyse the measures - results in Alarm sheet

## Original data
File name |  | HCMBB_A001-02000013_cc.xls
Component ID |  | HCMBB_A001 |  |  | Serial Number |  |  | 2000013 | [11 empty] | Signs
Ref. Test Proced. |  | CERN IT 2708/LHC/LHC Rev 1.1 Annex b.18 | [16 empty] | b even | 1
Firm Name |  | Ansaldo | [9 empty] | File name for aperture 1 |  |  | ansaldo13cct1 |  |  |  | a odd | 1
Test Type |  | Magnetic Measures |  |  | ITP number |  |  | 21 |  |  |  | File name for aperture 2 |  |  | ansaldo13cct2 |  |  |  | a even | 1
Test Operator |  | G. Peiro - A. Musso |  |  | Test Controller |  |  | E. Todesco
Aperture 1 | [11 empty] | Aperture 2
Date of test |  | 2002-06-26 |  |  | Time of test |  |  | 10.00-13.00 |  |  |  | Date of test |  | 2002-06-26 |  |  | Time of test |  |  | 14.00-17.00
Mole name |  | Dipole 5 |  |  | Calibration date |  |  | 20.07.99 |  |  |  | Mole name |  | Dipole 5 |  |  | Calibration date |  |  | 20.07.99
Coil length |  | 0.7499 m |  |  | Coil temperature |  |  | 27.5 |  |  |  | Coil length |  | 0.7499 m |  |  | Coil temperature |  |  | 29.1
Excitation current (A) |  | 8.5 |  |  | Analysis tools |  |  | dri,rot,nor,cel,fdw |  |  |  | Excitation current (A) |  | 8.5 |  |  | Analysis tools |  |  | dri,rot,nor,cel,fdw
N. points |  | 15 |  |  | N. measure |  |  | 10 (5+,5-) |  |  |  | N. points |  | 15 |  |  | N. measure |  |  | 10 (5+,5-)
Rotation speed |  | 0.6 |  |  | I. Gain |  |  | 50/500 |  |  |  | Rotation speed |  | 0.6 |  |  | I. Gain |  |  | 50/500
Dx average |  | 0.00019483 |  |  | Dy average |  |  | 0.001057 |  |  |  | Dx average |  | 0.00015213 |  |  | Dy average |  |  | 0.0010143
Magnetic length |  | 14.4427 |  |  | Tranfs. Function (mT/KA) |  |  | 594.6013 |  |  |  | Magnetic length |  | 14.4468 |  |  | Tranfs. Function (mT/KA) |  |  | 594.6209
Phase |  | -11.0294 |  |  | Integral field (T.m) |  |  | 0.072996 |  |  |  | Phase |  | -11.0075 |  |  | Integral field (T.m) |  |  | 0.07302
File | Aperture 1 - Collared coils | [19 empty] | Integrals | File | Aperture 2 - Collared coils | [19 empty] | Integrals
C1 (T) | 0.003241 | 0.005057 | 0.005054 | 0.005053 | 0.005054 | 0.005053 | 0.005053 | 0.005052 | 0.005052 | 0.005054 | 0.005054 | 0.005054 | 0.005054 | 0.005053 | 0.005055 | 0.005055 | 0.005055 | 0.005055 | 0.005057 | 0.003125 | 0.072996 | C1 (T) | 0.003271 | 0.005057 | 0.005054 | 0.005054 | 0.005053 | 0.005054 | 0.005053 | 0.005053 | 0.005053 | 0.005053 | 0.005053 | 0.005053 | 0.005053 | 0.005054 | 0.005056 | 0.005056 | 0.005056 | 0.005055 | 0.005057 | 0.003123 | 0.07302
Angle (mrad) | 4.8215 | -0.85876 | -0.70298 | -0.26237 | 0.31758 | 0.083572 | -0.42302 | -0.96274 | -0.62667 | 0.21082 | 0.57278 | 0.078234 | 0.44029 | 0.70748 | 0.47942 | -0.25937 | -0.30614 | -0.2946 | -0.62144 | -1.0619 | -11.0294 | Angle (mrad) | 4.3328 | -0.68588 | -0.46871 | -0.017209 | 0.81563 | 0.10181 | -0.50004 | -0.3785 | 0.060913 | 0.34287 | 0.45034 | 0.11184 | 0.074246 | 0.074946 | 0.064183 | -0.29601 | -0.50895 | -0.45078 | -0.39042 | -1.7589 | -11.0075
Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20 |  | Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20
b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000
b2 | -4.0906 | -0.47182 | -1.4612 | -1.1782 | 0.14346 | 0.25919 | 0.65373 | 0.78724 | -0.00045971 | -0.55722 | -0.12767 | 0.71366 | -0.38772 | -1.082 | -0.5893 | 0.018601 | 0.2513 | -0.1881 | -0.090753 | -0.11639 | -0.31165 | b2 | -3.9627 | 1.0408 | 1.2976 | 0.24246 | 1.0909 | 0.82241 | 1.3601 | 1.1621 | -0.20553 | -0.49445 | -0.17838 | 0.3022 | 0.27175 | -0.46188 | 0.73615 | 0.24661 | -0.44898 | -0.57962 | -0.27071 | 0.9101 | 0.20407
b3 | 16.6984 | -4.5214 | -4.9797 | -4.8992 | -4.9373 | -4.4646 | -4.9832 | -4.9152 | -5.0293 | -4.9093 | -4.2337 | -4.3964 | -4.4649 | -4.7123 | -5.0862 | -5.1368 | -4.9379 | -5.2304 | -4.6789 | -12.4569 | -4.3362 | b3 | 17.4317 | -4.5963 | -4.6239 | -4.107 | -4.6889 | -4.8376 | -4.7625 | -4.6397 | -4.329 | -4.2595 | -4.3152 | -4.1654 | -3.8146 | -4.0985 | -4.3163 | -3.9454 | -4.0534 | -4.6568 | -4.6117 | -11.3908 | -3.8712
b4 | -2.1333 | -0.096291 | 0.023548 | -0.028886 | -0.043841 | 0.041818 | -0.025003 | -0.089603 | -0.11177 | -0.038883 | 0.017487 | -0.16483 | -0.11078 | -0.1501 | -0.1566 | 0.036252 | -0.21663 | -0.31421 | -0.010666 | -1.0901 | -0.18073 | b4 | -0.83983 | 0.016734 | 0.23327 | 0.14193 | -0.039227 | -0.06637 | -0.023905 | -0.17636 | -0.11842 | -0.084638 | 0.058226 | -0.084785 | -0.009589 | -0.11355 | -0.19515 | 0.096198 | -0.014211 | 0.013042 | 0.10101 | 0.27454 | -0.033194
b5 | -7.3864 | 0.020762 | 0.10729 | 0.16946 | 0.11205 | 0.247 | 0.45738 | 0.43426 | 0.34348 | 0.26937 | 0.30092 | 0.26805 | 0.32749 | 0.13786 | 0.2146 | 0.12051 | 0.11648 | 0.12344 | 0.013759 | -4.9933 | -0.20974 | b5 | -7.1816 | 0.24326 | 0.062391 | 0.25528 | 0.35498 | 0.57729 | 0.57987 | 0.48171 | 0.46842 | 0.51789 | 0.51743 | 0.44275 | 0.4644 | 0.486 | 0.52616 | 0.46845 | 0.3397 | 0.34172 | 0.20191 | -4.9477 | -0.019473
b6 | -0.057414 | -0.010084 | -0.098868 | -0.020086 | -0.069775 | 0.068348 | 0.063962 | 0.15444 | 0.024497 | -0.056464 | -0.11394 | -0.072606 | -0.0055783 | -0.044327 | -0.036624 | 0.0036246 | -0.030522 | 0.012374 | 0.037299 | 0.01799 | -0.011427 | b6 | -0.29512 | 0.008176 | 0.031822 | 0.08788 | 0.014384 | -0.096437 | -0.0016733 | 0.0054139 | 0.036078 | -0.071439 | 0.072072 | -0.099088 | -0.032953 | -0.19465 | -0.018149 | -0.065584 | -0.016036 | -0.095257 | 0.01467 | -0.0047589 | -0.031906
b7 | 1.5294 | 0.79174 | 0.90875 | 0.92401 | 0.92432 | 0.90677 | 0.97242 | 0.95736 | 0.95794 | 0.91051 | 0.95416 | 0.94682 | 0.96535 | 0.93106 | 0.89568 | 0.91009 | 0.90166 | 0.91939 | 0.87913 | 0.10193 | 0.91387 | b7 | 1.6539 | 0.86552 | 0.90843 | 0.93595 | 0.95239 | 0.90228 | 1.0206 | 1.0066 | 1.0048 | 0.95574 | 0.93887 | 1.0079 | 0.98282 | 1.0165 | 0.90933 | 0.98047 | 0.92691 | 0.96118 | 0.94614 | 0.13699 | 0.95392
b8 | -0.070051 | 0.036249 | -0.019181 | 0.039894 | 0.0096155 | 0.021715 | -0.033095 | 0.028389 | -0.0030265 | -0.03764 | -0.045463 | -0.049247 | 0.034902 | -0.014668 | 0.0028872 | -0.014223 | 0.00033595 | -0.04601 | 0.014139 | 0.12837 | -0.0020741 | b8 | -0.00076641 | -0.014217 | -0.033802 | 0.060089 | 0.02376 | -0.050652 | 0.029687 | 0.013022 | 0.025735 | -0.028335 | 0.05244 | -0.05276 | -0.0037441 | -0.05679 | -0.00051635 | -0.038504 | 0.048073 | -0.038083 | 0.0054846 | -0.024742 | -0.003887
b9 | 0.2808 | 0.56475 | 0.56467 | 0.56871 | 0.57378 | 0.58759 | 0.58344 | 0.5866 | 0.56986 | 0.56848 | 0.58408 | 0.58109 | 0.59056 | 0.57365 | 0.5777 | 0.57697 | 0.57619 | 0.57052 | 0.57365 | 0.35211 | 0.55921 | b9 | 0.21848 | 0.57431 | 0.56779 | 0.58562 | 0.58475 | 0.57504 | 0.5858 | 0.57953 | 0.58099 | 0.57521 | 0.58875 | 0.58669 | 0.58382 | 0.5787 | 0.57162 | 0.58335 | 0.58623 | 0.57764 | 0.57539 | 0.34177 | 0.56028
b10 | 0.063706 | 0.034618 | -0.03713 | 0.023164 | -0.027842 | 0.044007 | -0.034735 | 0.092535 | -0.016258 | -0.068483 | -0.055056 | -0.026148 | 0.096827 | -0.022508 | -0.0021577 | -0.027721 | 0.032992 | -0.030131 | -0.0046276 | 0.15841 | 0.0057185 | b10 | -0.023583 | 0.038587 | -0.041186 | 0.11288 | 0.029064 | -0.071481 | 0.065687 | 0.053532 | 0.051322 | -0.055742 | 0.098154 | -0.066179 | 0.025706 | -0.080954 | 0.0053174 | -0.068216 | 0.082547 | -0.066887 | -0.0070381 | -0.034806 | 0.0035473
b11 | 0.60066 | 0.74289 | 0.73941 | 0.73732 | 0.74103 | 0.73981 | 0.74026 | 0.73325 | 0.73422 | 0.74113 | 0.73734 | 0.74085 | 0.74232 | 0.7373 | 0.74689 | 0.74715 | 0.74294 | 0.74061 | 0.74095 | 0.58567 | 0.7307 | b11 | 0.58855 | 0.73707 | 0.73553 | 0.74053 | 0.74232 | 0.74184 | 0.74312 | 0.7373 | 0.73403 | 0.73882 | 0.74159 | 0.73486 | 0.73806 | 0.73634 | 0.75482 | 0.74936 | 0.75246 | 0.74158 | 0.73611 | 0.57509 | 0.73044
b12 | 0.004028 | 0.0084746 | -0.0068864 | 0.0012303 | -0.0034212 | 0.0048615 | -0.0066884 | 0.010809 | -0.0033986 | -0.013268 | -0.0085776 | -0.0035969 | 0.012812 | -0.0026165 | -0.0016929 | -0.0062195 | 0.001758 | -0.0054657 | 0.00052149 | 0.012439 | -0.00057576 | b12 | 0.011344 | 0.0039764 | -0.0069948 | 0.016282 | 0.006516 | -0.01038 | 0.011651 | 0.009925 | 0.0059275 | -0.0062462 | 0.013074 | -0.0063418 | 0.0048835 | -0.0071068 | 0.0056318 | -0.009131 | 0.013962 | -0.0099757 | -0.0029067 | -0.0056653 | 0.001899
b13 | 0.071989 | 0.082114 | 0.083407 | 0.082494 | 0.084549 | 0.080956 | 0.080201 | 0.080209 | 0.080466 | 0.080407 | 0.082508 | 0.080854 | 0.07792 | 0.081965 | 0.083284 | 0.08316 | 0.081297 | 0.080572 | 0.084183 | 0.039955 | 0.080034 | b13 | 0.069848 | 0.07995 | 0.079838 | 0.08375 | 0.078322 | 0.08087 | 0.078288 | 0.078856 | 0.080008 | 0.079817 | 0.081745 | 0.081741 | 0.079306 | 0.081845 | 0.083328 | 0.083524 | 0.084497 | 0.078395 | 0.08032 | 0.045883 | 0.079312
b14 | -0.0050729 | -0.0084342 | -0.013154 | -0.0096047 | -0.011973 | -0.0058324 | -0.014029 | -0.0016153 | -0.012009 | -0.018373 | -0.014502 | -0.01278 | -0.0019815 | -0.01168 | -0.011264 | -0.013624 | -0.0073361 | -0.012302 | -0.0087308 | -0.0024662 | -0.010073 | b14 | 0.0059332 | -0.0060342 | -0.011106 | 0.0010795 | -0.0036019 | -0.012629 | -0.002859 | -0.0024205 | -0.0025062 | -0.011923 | 0.0018494 | -0.013095 | -0.0049672 | -0.012796 | -0.0051931 | -0.012752 | -0.0010951 | -0.016389 | -0.010271 | -0.0086721 | -0.0066579
b15 | 0.012743 | 0.040838 | 0.040229 | 0.048208 | 0.04198 | 0.046987 | 0.047569 | 0.046971 | 0.048518 | 0.052575 | 0.046305 | 0.048048 | 0.050394 | 0.046199 | 0.050498 | 0.04672 | 0.046928 | 0.047756 | 0.039057 | 0.018475 | 0.044413 | b15 | 0.0051604 | 0.049489 | 0.044595 | 0.042263 | 0.051758 | 0.043406 | 0.053205 | 0.04567 | 0.050213 | 0.04906 | 0.042237 | 0.044907 | 0.047927 | 0.040244 | 0.044454 | 0.0414 | 0.043605 | 0.05182 | 0.049682 | 0.032659 | 0.044611
b16 | [21 empty] | b16
b17 | [21 empty] | b17
a1 | 48.2157 | -8.5876 | -7.0298 | -2.6237 | 3.1758 | 0.83572 | -4.2302 | -9.6274 | -6.2667 | 2.1082 | 5.7278 | 0.78234 | 4.403 | 7.0748 | 4.7942 | -2.5937 | -3.0614 | -2.946 | -6.2144 | -10.619 | -1.4853e-14 | a1 | 43.3279 | -6.8588 | -4.6871 | -0.17209 | 8.1563 | 1.0181 | -5.0004 | -3.785 | 0.60913 | 3.4287 | 4.5034 | 1.1184 | 0.74246 | 0.74946 | 0.64183 | -2.9601 | -5.0895 | -4.5078 | -3.9042 | -17.5885 | 0
a2 | 10.452 | -1.2651 | -1.3394 | -1.2319 | -1.1201 | -1.0994 | -0.57231 | -0.48797 | -1.3866 | -0.89415 | -1.1266 | -2.6571 | -1.7011 | -1.412 | -1.5414 | -1.5245 | -1.6665 | -0.85805 | -0.94554 | -9.2324 | -1.1339 | a2 | 6.5221 | 1.1272 | 1.0369 | 1.2911 | 0.92183 | 1.8649 | 1.3715 | 0.54566 | 0.68174 | 1.0026 | 0.56599 | 0.63839 | 0.86327 | 0.74057 | 0.71416 | 0.50525 | 0.083036 | -0.02758 | 0.50019 | -3.6779 | 0.84994
a3 | -3.5169 | -1.2414 | -0.19329 | -0.77057 | -0.42387 | -0.54027 | -0.085493 | -0.14649 | -0.33577 | -0.65991 | -0.16204 | -0.20959 | -0.12693 | -0.65216 | -0.26048 | -0.34109 | -0.46558 | -0.27746 | -0.57083 | -0.094723 | -0.50765 | a3 | -1.6778 | -0.38618 | 0.072068 | -0.065414 | 0.37737 | -0.0076867 | -0.42962 | -0.55847 | -0.069521 | -0.88853 | 0.096686 | -0.21758 | -0.18689 | 0.11035 | -0.010481 | -0.25828 | -0.29953 | -0.42491 | -0.57276 | 0.31898 | -0.23914
a4 | 2.4934 | 0.038661 | -0.11495 | -0.13217 | -0.10001 | -0.16155 | -0.05561 | 0.13286 | -0.014816 | -0.46769 | -0.35321 | -0.22067 | -0.25977 | -0.22285 | -0.099853 | -0.12583 | -0.13919 | -0.14291 | 0.00030551 | -1.0106 | -0.076091 | a4 | 1.2381 | -0.13933 | -0.35792 | 0.010514 | 0.031637 | 0.15856 | -0.11793 | -0.072412 | 0.22023 | 0.089457 | 0.13826 | 0.23355 | 0.19071 | 0.11225 | 0.021158 | -0.15521 | -0.20834 | -0.27145 | 0.020798 | 0.2744 | 0.045413
a5 | 1.6315 | -0.43322 | -0.23405 | -0.24758 | -0.074624 | -0.087097 | 0.036429 | 0.14353 | 0.082908 | -0.15438 | -0.096067 | -0.1959 | -0.16723 | -0.19381 | 0.058045 | -0.16112 | -0.2289 | -0.091156 | -0.1909 | -0.082814 | -0.064425 | a5 | 2.0464 | -0.26848 | -0.038237 | -0.061186 | 0.080057 | -0.062896 | -0.046328 | 0.047811 | 0.047542 | -0.10822 | -0.025806 | -0.092434 | -0.067994 | 0.023314 | 0.15979 | -0.028912 | -0.0087675 | -0.013145 | -0.064479 | -0.048267 | 0.039762
a6 | 0.39563 | 0.055054 | -0.013801 | -0.025993 | 0.023477 | -0.0071092 | 0.025983 | 0.061751 | -0.077091 | -0.094542 | -0.039825 | -0.042274 | -0.066133 | -0.02823 | -0.006404 | -0.029465 | -0.014442 | -0.075952 | -0.0049251 | -0.29254 | -0.014908 | a6 | -0.061993 | -0.020899 | -0.052406 | 0.098881 | -0.0018421 | 0.089479 | -0.043302 | 0.02463 | 0.085883 | 0.091105 | 0.027939 | 0.052138 | -0.0069503 | 0.050253 | -0.0091117 | 0.045701 | 0.024152 | -0.032426 | 0.0097303 | -0.15684 | 0.015358
a7 | 1.3603 | -0.041012 | -0.037059 | -0.000299 | -0.013264 | 0.047369 | 0.064089 | 0.049624 | 0.037028 | -0.0091815 | -0.059179 | -0.053027 | 0.0046108 | -0.018821 | 0.072475 | 0.011783 | -0.020688 | 0.03555 | -0.054973 | 0.16916 | 0.051487 | a7 | 1.5727 | -0.068245 | -0.017133 | -0.018997 | 0.037345 | 0.060246 | 0.049725 | 0.040226 | 0.14756 | 0.051664 | 0.029881 | -0.0026404 | -0.0055311 | 0.016548 | 0.1215 | 0.023817 | 0.022179 | 0.035575 | -0.060259 | 0.051442 | 0.078492
a8 | 0.0076448 | 0.012557 | -0.01158 | -0.017008 | 0.016752 | 0.023972 | 0.012853 | 0.053586 | -0.0044769 | -0.02257 | -0.0049429 | 0.014389 | -0.0095054 | 0.02396 | 0.011514 | -0.0054928 | 0.017834 | -0.018181 | -0.0076205 | -0.010571 | 0.0043815 | a8 | -0.015878 | 0.0296 | 0.012787 | 0.092354 | 0.028707 | 0.063166 | 0.00034165 | 0.0053198 | 0.027737 | 0.027423 | 0.043489 | 0.013216 | 0.017562 | 0.053792 | 0.0321 | 0.040429 | 0.043757 | -0.023519 | -0.0055277 | -0.017324 | 0.025006
a9 | -0.32776 | -0.081106 | -0.050597 | -0.050368 | -0.039677 | -0.0089274 | -0.0042191 | 0.0075199 | -0.02248 | -0.032918 | 0.011581 | -0.036658 | -0.028357 | -0.027822 | 0.0036364 | -0.012509 | -0.029183 | -0.027822 | -0.036735 | -0.069645 | -0.03738 | a9 | -0.2666 | -0.039023 | -0.01652 | -0.031135 | -0.036643 | 0.0145 | -0.0069591 | 0.0087856 | -0.018904 | -0.016447 | 0.016321 | -0.024476 | -0.013837 | -0.037993 | 0.0060744 | -0.012309 | -0.00028437 | -0.015414 | -0.0367 | -0.041126 | -0.023852
a10 | 0.051749 | 0.076593 | 0.045855 | -0.00044719 | 0.035256 | 0.013523 | -0.0078405 | 0.018582 | -0.011873 | -0.062663 | -0.0099946 | -0.01936 | -0.018478 | 0.0064355 | -0.0034752 | -0.0007086 | 0.013073 | -0.018115 | 0.044153 | 0.018619 | 0.0075436 | a10 | -0.035499 | 0.0011601 | -0.0080573 | 0.066427 | -0.022137 | 0.02566 | -0.033617 | 0.02611 | -0.01054 | -0.006493 | 0.050199 | -0.0038367 | -0.0068444 | 0.034238 | 0.03117 | 0.052431 | 0.057177 | -0.05359 | -0.02384 | -0.046036 | 0.0064474
a11 | 0.13482 | -0.063561 | -0.065303 | -0.057247 | -0.048854 | -0.040926 | -0.040105 | -0.042486 | -0.041536 | -0.047936 | -0.045457 | -0.054911 | -0.046178 | -0.041782 | -0.033727 | -0.048218 | -0.05166 | -0.046312 | -0.057186 | -0.024144 | -0.041639 | a11 | 0.13708 | -0.066396 | -0.053563 | -0.047131 | -0.039965 | -0.033856 | -0.04081 | -0.039058 | -0.028334 | -0.03795 | -0.036665 | -0.045789 | -0.045235 | -0.044002 | -0.03245 | -0.041289 | -0.042216 | -0.047853 | -0.056995 | -0.033023 | -0.036969
a12 | -0.010006 | 0.0087107 | 0.001254 | -0.0043509 | 0.0033887 | -0.0015421 | -0.0010289 | 0.00090475 | -0.0071093 | -0.011804 | -0.0032323 | -0.0031488 | -0.0081317 | -0.003855 | -0.0045524 | -0.0043189 | -0.003162 | -0.0097388 | 0.00084307 | -0.01095 | -0.0033257 | a12 | -0.0050551 | -0.0019499 | -0.0023327 | 0.0095352 | -0.0034138 | 0.0063355 | -0.0071689 | 0.0018115 | -0.00096008 | 0.0029091 | 0.0060441 | -0.00015281 | -0.003149 | 0.0041533 | 0.00013366 | 0.0061318 | 0.0059066 | -0.0088014 | -0.0019131 | -0.0039296 | 0.00038497
a13 | -0.023261 | -0.013971 | -0.0083952 | -0.0098542 | -0.0058263 | -0.0061481 | -0.0037788 | -0.0055518 | -0.0061909 | -0.0058886 | -0.0033123 | -0.0081124 | -0.0079374 | -0.0068017 | -0.0034483 | -0.0055854 | -0.0085047 | -0.0068895 | -0.008748 | -0.0025784 | -0.0073452 | a13 | -0.034786 | -0.0094755 | -0.0049248 | -0.0094703 | -0.0060521 | -0.0022242 | -0.005812 | -0.0059621 | -0.0057747 | -0.0044744 | -0.0053168 | -0.0054259 | -0.0040595 | -0.0042799 | -0.0032507 | -0.0034959 | -0.0073414 | -0.0035822 | -0.0079171 | -0.0046385 | -0.0064542
a14 | 0.0097799 | -0.018438 | -0.018986 | -0.024228 | -0.020012 | -0.02091 | -0.02143 | -0.020762 | -0.023278 | -0.027118 | -0.02256 | -0.025233 | -0.025453 | -0.022146 | -0.023949 | -0.022867 | -0.021715 | -0.023189 | -0.017548 | -0.00045497 | -0.020448 | a14 | 0.0033077 | -0.021955 | -0.023354 | -0.01693 | -0.026289 | -0.018163 | -0.027377 | -0.021887 | -0.023693 | -0.021914 | -0.019043 | -0.020378 | -0.021323 | -0.016626 | -0.019672 | -0.018818 | -0.019974 | -0.02638 | -0.023846 | -0.0057246 | -0.020192
a15 | 0.0086944 | -0.012886 | -0.018288 | -0.013299 | -0.018267 | -0.010556 | -0.021836 | -0.0056397 | -0.018863 | -0.026744 | -0.022627 | -0.019928 | -0.0049733 | -0.018688 | -0.016012 | -0.019965 | -0.012324 | -0.020456 | -0.015747 | 0.011409 | -0.014771 | a15 | -0.0078717 | -0.011547 | -0.020163 | -0.001782 | -0.0084071 | -0.020917 | -0.0068684 | -0.007355 | -0.0065404 | -0.020329 | -0.00093479 | -0.021237 | -0.011109 | -0.020845 | -0.010673 | -0.019171 | -0.0035188 | -0.02302 | -0.015735 | -0.01741 | -0.012774
a16 | [21 empty] | a16
a17 | [21 empty] | a17
Dx (m) | -0.00020295 | -6.3757e-05 | 9.3946e-05 | -5.3168e-05 | 6.8905e-05 | -9.9102e-05 | 7.8562e-05 | -0.00021133 | 3.5973e-05 | 0.00014717 | 0.00012504 | 5.6398e-05 | -0.00022351 | 5.257e-05 | 4.5448e-06 | 6.2709e-05 | -7.3059e-05 | 6.6303e-05 | 1.8327e-05 | -0.00045682 | 0.00019483 | Dx (m) | 8.7266e-05 | -8.8043e-05 | 9.334e-05 | -0.00024843 | -6.909e-05 | 0.00016614 | -0.00015403 | -0.00011991 | -0.00011963 | 0.00012716 | -0.00021878 | 0.00015195 | -5.995e-05 | 0.00019094 | -8.9414e-06 | 0.00016082 | -0.00017868 | 0.0001448 | 1.192e-05 | 9.4761e-05 | 0.00015213
Dy (m) | -0.00010091 | -0.00018073 | -9.713e-05 | -3.097e-06 | -7.6337e-05 | -3.6556e-05 | 2.2262e-05 | -5.5327e-05 | 2.9526e-05 | 0.00015326 | 3.0752e-05 | 4.8605e-05 | 2.8412e-05 | -1.1859e-05 | 8.1151e-06 | 5.6593e-06 | -3.4994e-05 | 4.5727e-05 | -9.9888e-05 | -7.2876e-05 | 0.001057 | Dy (m) | 8.221e-05 | -1.0607e-05 | 2.542e-05 | -0.00016831 | 4.6977e-05 | -5.1219e-05 | 6.8446e-05 | -6.6553e-05 | 1.9793e-05 | 2.1472e-05 | -0.00012589 | 1.8344e-05 | 1.2091e-05 | -6.7635e-05 | -7.0585e-05 | -0.00011008 | -0.0001392 | 0.00013219 | 5.5978e-05 | 0.00014153 | 0.0010143
Magnetic length (m) |  | 14.4427 | [19 empty] | Magnetic length (m) |  | 14.4468
Transfer funct. (mT/KA) |  | 594.6013 | [19 empty] | Transfer funct. (mT/KA) |  | 594.6209
Coil positioning (mm) |  | 0.016382 | [19 empty] | Coil positioning (mm) |  | 0.015878
Current (A) |  | 8.5 | [19 empty] | Current (A) |  | 8.5

## Summary Data
File | HCMBB_A001-2000013 |  |  | Aperture 1 - Collared coils | [16 empty] | Integrals | File | HCMBB_A001-2000013 |  |  | Aperture 2 - Collared coils | [16 empty] | Integrals
C1/i (mT/KA) | 381.2941 | 594.9412 | 594.5882 | 594.4706 | 594.5882 | 594.4706 | 594.4706 | 594.3529 | 594.3529 | 594.5882 | 594.5882 | 594.5882 | 594.5882 | 594.4706 | 594.7059 | 594.7059 | 594.7059 | 594.7059 | 594.9412 | 367.6471 | 0.072996 | C1/i (mT/KA) | 384.8235 | 594.9412 | 594.5882 | 594.5882 | 594.4706 | 594.5882 | 594.4706 | 594.4706 | 594.4706 | 594.4706 | 594.4706 | 594.4706 | 594.4706 | 594.5882 | 594.8235 | 594.8235 | 594.8235 | 594.7059 | 594.9412 | 367.4118 | 0.07302
Angle (mrad) | 4.8215 | -0.85876 | -0.70298 | -0.26237 | 0.31758 | 0.083572 | -0.42302 | -0.96274 | -0.62667 | 0.21082 | 0.57278 | 0.078234 | 0.44029 | 0.70748 | 0.47942 | -0.25937 | -0.30614 | -0.2946 | -0.62144 | -1.0619 | -11.0294 | Angle (mrad) | 4.3328 | -0.68588 | -0.46871 | -0.017209 | 0.81563 | 0.10181 | -0.50004 | -0.3785 | 0.060913 | 0.34287 | 0.45034 | 0.11184 | 0.074246 | 0.074946 | 0.064183 | -0.29601 | -0.50895 | -0.45078 | -0.39042 | -1.7589 | -11.0075
Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20 | 0 | Multipoles | Position 1 | Position 2 | Position 3 | Position 4 | Position 5 | Position 6 | Position 7 | Position 8 | Position 9 | Position 10 | Position 11 | Position 12 | Position 13 | Position 14 | Position 15 | Position 16 | Position 17 | Position 18 | Position 19 | Position 20 | 0
b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | b1 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000
b2 | -4.0906 | -0.47182 | -1.4612 | -1.1782 | 0.14346 | 0.25919 | 0.65373 | 0.78724 | -0.00045971 | -0.55722 | -0.12767 | 0.71366 | -0.38772 | -1.082 | -0.5893 | 0.018601 | 0.2513 | -0.1881 | -0.090753 | -0.11639 | -0.31165 | b2 | -3.9627 | 1.0408 | 1.2976 | 0.24246 | 1.0909 | 0.82241 | 1.3601 | 1.1621 | -0.20553 | -0.49445 | -0.17838 | 0.3022 | 0.27175 | -0.46188 | 0.73615 | 0.24661 | -0.44898 | -0.57962 | -0.27071 | 0.9101 | 0.20407
b3 | 16.6984 | -4.5214 | -4.9797 | -4.8992 | -4.9373 | -4.4646 | -4.9832 | -4.9152 | -5.0293 | -4.9093 | -4.2337 | -4.3964 | -4.4649 | -4.7123 | -5.0862 | -5.1368 | -4.9379 | -5.2304 | -4.6789 | -12.4569 | -4.3362 | b3 | 17.4317 | -4.5963 | -4.6239 | -4.107 | -4.6889 | -4.8376 | -4.7625 | -4.6397 | -4.329 | -4.2595 | -4.3152 | -4.1654 | -3.8146 | -4.0985 | -4.3163 | -3.9454 | -4.0534 | -4.6568 | -4.6117 | -11.3908 | -3.8712
b4 | -2.1333 | -0.096291 | 0.023548 | -0.028886 | -0.043841 | 0.041818 | -0.025003 | -0.089603 | -0.11177 | -0.038883 | 0.017487 | -0.16483 | -0.11078 | -0.1501 | -0.1566 | 0.036252 | -0.21663 | -0.31421 | -0.010666 | -1.0901 | -0.18073 | b4 | -0.83983 | 0.016734 | 0.23327 | 0.14193 | -0.039227 | -0.06637 | -0.023905 | -0.17636 | -0.11842 | -0.084638 | 0.058226 | -0.084785 | -0.009589 | -0.11355 | -0.19515 | 0.096198 | -0.014211 | 0.013042 | 0.10101 | 0.27454 | -0.033194
b5 | -7.3864 | 0.020762 | 0.10729 | 0.16946 | 0.11205 | 0.247 | 0.45738 | 0.43426 | 0.34348 | 0.26937 | 0.30092 | 0.26805 | 0.32749 | 0.13786 | 0.2146 | 0.12051 | 0.11648 | 0.12344 | 0.013759 | -4.9933 | -0.20974 | b5 | -7.1816 | 0.24326 | 0.062391 | 0.25528 | 0.35498 | 0.57729 | 0.57987 | 0.48171 | 0.46842 | 0.51789 | 0.51743 | 0.44275 | 0.4644 | 0.486 | 0.52616 | 0.46845 | 0.3397 | 0.34172 | 0.20191 | -4.9477 | -0.019473
b6 | -0.057414 | -0.010084 | -0.098868 | -0.020086 | -0.069775 | 0.068348 | 0.063962 | 0.15444 | 0.024497 | -0.056464 | -0.11394 | -0.072606 | -0.0055783 | -0.044327 | -0.036624 | 0.0036246 | -0.030522 | 0.012374 | 0.037299 | 0.01799 | -0.011427 | b6 | -0.29512 | 0.008176 | 0.031822 | 0.08788 | 0.014384 | -0.096437 | -0.0016733 | 0.0054139 | 0.036078 | -0.071439 | 0.072072 | -0.099088 | -0.032953 | -0.19465 | -0.018149 | -0.065584 | -0.016036 | -0.095257 | 0.01467 | -0.0047589 | -0.031906
b7 | 1.5294 | 0.79174 | 0.90875 | 0.92401 | 0.92432 | 0.90677 | 0.97242 | 0.95736 | 0.95794 | 0.91051 | 0.95416 | 0.94682 | 0.96535 | 0.93106 | 0.89568 | 0.91009 | 0.90166 | 0.91939 | 0.87913 | 0.10193 | 0.91387 | b7 | 1.6539 | 0.86552 | 0.90843 | 0.93595 | 0.95239 | 0.90228 | 1.0206 | 1.0066 | 1.0048 | 0.95574 | 0.93887 | 1.0079 | 0.98282 | 1.0165 | 0.90933 | 0.98047 | 0.92691 | 0.96118 | 0.94614 | 0.13699 | 0.95392
b8 | -0.070051 | 0.036249 | -0.019181 | 0.039894 | 0.0096155 | 0.021715 | -0.033095 | 0.028389 | -0.0030265 | -0.03764 | -0.045463 | -0.049247 | 0.034902 | -0.014668 | 0.0028872 | -0.014223 | 0.00033595 | -0.04601 | 0.014139 | 0.12837 | -0.0020741 | b8 | -0.00076641 | -0.014217 | -0.033802 | 0.060089 | 0.02376 | -0.050652 | 0.029687 | 0.013022 | 0.025735 | -0.028335 | 0.05244 | -0.05276 | -0.0037441 | -0.05679 | -0.00051635 | -0.038504 | 0.048073 | -0.038083 | 0.0054846 | -0.024742 | -0.003887
b9 | 0.2808 | 0.56475 | 0.56467 | 0.56871 | 0.57378 | 0.58759 | 0.58344 | 0.5866 | 0.56986 | 0.56848 | 0.58408 | 0.58109 | 0.59056 | 0.57365 | 0.5777 | 0.57697 | 0.57619 | 0.57052 | 0.57365 | 0.35211 | 0.55921 | b9 | 0.21848 | 0.57431 | 0.56779 | 0.58562 | 0.58475 | 0.57504 | 0.5858 | 0.57953 | 0.58099 | 0.57521 | 0.58875 | 0.58669 | 0.58382 | 0.5787 | 0.57162 | 0.58335 | 0.58623 | 0.57764 | 0.57539 | 0.34177 | 0.56028
b10 | 0.063706 | 0.034618 | -0.03713 | 0.023164 | -0.027842 | 0.044007 | -0.034735 | 0.092535 | -0.016258 | -0.068483 | -0.055056 | -0.026148 | 0.096827 | -0.022508 | -0.0021577 | -0.027721 | 0.032992 | -0.030131 | -0.0046276 | 0.15841 | 0.0057185 | b10 | -0.023583 | 0.038587 | -0.041186 | 0.11288 | 0.029064 | -0.071481 | 0.065687 | 0.053532 | 0.051322 | -0.055742 | 0.098154 | -0.066179 | 0.025706 | -0.080954 | 0.0053174 | -0.068216 | 0.082547 | -0.066887 | -0.0070381 | -0.034806 | 0.0035473
b11 | 0.60066 | 0.74289 | 0.73941 | 0.73732 | 0.74103 | 0.73981 | 0.74026 | 0.73325 | 0.73422 | 0.74113 | 0.73734 | 0.74085 | 0.74232 | 0.7373 | 0.74689 | 0.74715 | 0.74294 | 0.74061 | 0.74095 | 0.58567 | 0.7307 | b11 | 0.58855 | 0.73707 | 0.73553 | 0.74053 | 0.74232 | 0.74184 | 0.74312 | 0.7373 | 0.73403 | 0.73882 | 0.74159 | 0.73486 | 0.73806 | 0.73634 | 0.75482 | 0.74936 | 0.75246 | 0.74158 | 0.73611 | 0.57509 | 0.73044
b12 | 0.004028 | 0.0084746 | -0.0068864 | 0.0012303 | -0.0034212 | 0.0048615 | -0.0066884 | 0.010809 | -0.0033986 | -0.013268 | -0.0085776 | -0.0035969 | 0.012812 | -0.0026165 | -0.0016929 | -0.0062195 | 0.001758 | -0.0054657 | 0.00052149 | 0.012439 | -0.00057576 | b12 | 0.011344 | 0.0039764 | -0.0069948 | 0.016282 | 0.006516 | -0.01038 | 0.011651 | 0.009925 | 0.0059275 | -0.0062462 | 0.013074 | -0.0063418 | 0.0048835 | -0.0071068 | 0.0056318 | -0.009131 | 0.013962 | -0.0099757 | -0.0029067 | -0.0056653 | 0.001899
b13 | 0.071989 | 0.082114 | 0.083407 | 0.082494 | 0.084549 | 0.080956 | 0.080201 | 0.080209 | 0.080466 | 0.080407 | 0.082508 | 0.080854 | 0.07792 | 0.081965 | 0.083284 | 0.08316 | 0.081297 | 0.080572 | 0.084183 | 0.039955 | 0.080034 | b13 | 0.069848 | 0.07995 | 0.079838 | 0.08375 | 0.078322 | 0.08087 | 0.078288 | 0.078856 | 0.080008 | 0.079817 | 0.081745 | 0.081741 | 0.079306 | 0.081845 | 0.083328 | 0.083524 | 0.084497 | 0.078395 | 0.08032 | 0.045883 | 0.079312
b14 | -0.0050729 | -0.0084342 | -0.013154 | -0.0096047 | -0.011973 | -0.0058324 | -0.014029 | -0.0016153 | -0.012009 | -0.018373 | -0.014502 | -0.01278 | -0.0019815 | -0.01168 | -0.011264 | -0.013624 | -0.0073361 | -0.012302 | -0.0087308 | -0.0024662 | -0.010073 | b14 | 0.0059332 | -0.0060342 | -0.011106 | 0.0010795 | -0.0036019 | -0.012629 | -0.002859 | -0.0024205 | -0.0025062 | -0.011923 | 0.0018494 | -0.013095 | -0.0049672 | -0.012796 | -0.0051931 | -0.012752 | -0.0010951 | -0.016389 | -0.010271 | -0.0086721 | -0.0066579
b15 | 0.012743 | 0.040838 | 0.040229 | 0.048208 | 0.04198 | 0.046987 | 0.047569 | 0.046971 | 0.048518 | 0.052575 | 0.046305 | 0.048048 | 0.050394 | 0.046199 | 0.050498 | 0.04672 | 0.046928 | 0.047756 | 0.039057 | 0.018475 | 0.044413 | b15 | 0.0051604 | 0.049489 | 0.044595 | 0.042263 | 0.051758 | 0.043406 | 0.053205 | 0.04567 | 0.050213 | 0.04906 | 0.042237 | 0.044907 | 0.047927 | 0.040244 | 0.044454 | 0.0414 | 0.043605 | 0.05182 | 0.049682 | 0.032659 | 0.044611
b16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | b16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
b17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | b17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
a1 | 48.2157 | -8.5876 | -7.0298 | -2.6237 | 3.1758 | 0.83572 | -4.2302 | -9.6274 | -6.2667 | 2.1082 | 5.7278 | 0.78234 | 4.403 | 7.0748 | 4.7942 | -2.5937 | -3.0614 | -2.946 | -6.2144 | -10.619 | -1.4853e-14 | a1 | 43.3279 | -6.8588 | -4.6871 | -0.17209 | 8.1563 | 1.0181 | -5.0004 | -3.785 | 0.60913 | 3.4287 | 4.5034 | 1.1184 | 0.74246 | 0.74946 | 0.64183 | -2.9601 | -5.0895 | -4.5078 | -3.9042 | -17.5885 | 0
a2 | 10.452 | -1.2651 | -1.3394 | -1.2319 | -1.1201 | -1.0994 | -0.57231 | -0.48797 | -1.3866 | -0.89415 | -1.1266 | -2.6571 | -1.7011 | -1.412 | -1.5414 | -1.5245 | -1.6665 | -0.85805 | -0.94554 | -9.2324 | -1.1339 | a2 | 6.5221 | 1.1272 | 1.0369 | 1.2911 | 0.92183 | 1.8649 | 1.3715 | 0.54566 | 0.68174 | 1.0026 | 0.56599 | 0.63839 | 0.86327 | 0.74057 | 0.71416 | 0.50525 | 0.083036 | -0.02758 | 0.50019 | -3.6779 | 0.84994
a3 | -3.5169 | -1.2414 | -0.19329 | -0.77057 | -0.42387 | -0.54027 | -0.085493 | -0.14649 | -0.33577 | -0.65991 | -0.16204 | -0.20959 | -0.12693 | -0.65216 | -0.26048 | -0.34109 | -0.46558 | -0.27746 | -0.57083 | -0.094723 | -0.50765 | a3 | -1.6778 | -0.38618 | 0.072068 | -0.065414 | 0.37737 | -0.0076867 | -0.42962 | -0.55847 | -0.069521 | -0.88853 | 0.096686 | -0.21758 | -0.18689 | 0.11035 | -0.010481 | -0.25828 | -0.29953 | -0.42491 | -0.57276 | 0.31898 | -0.23914
a4 | 2.4934 | 0.038661 | -0.11495 | -0.13217 | -0.10001 | -0.16155 | -0.05561 | 0.13286 | -0.014816 | -0.46769 | -0.35321 | -0.22067 | -0.25977 | -0.22285 | -0.099853 | -0.12583 | -0.13919 | -0.14291 | 0.00030551 | -1.0106 | -0.076091 | a4 | 1.2381 | -0.13933 | -0.35792 | 0.010514 | 0.031637 | 0.15856 | -0.11793 | -0.072412 | 0.22023 | 0.089457 | 0.13826 | 0.23355 | 0.19071 | 0.11225 | 0.021158 | -0.15521 | -0.20834 | -0.27145 | 0.020798 | 0.2744 | 0.045413
a5 | 1.6315 | -0.43322 | -0.23405 | -0.24758 | -0.074624 | -0.087097 | 0.036429 | 0.14353 | 0.082908 | -0.15438 | -0.096067 | -0.1959 | -0.16723 | -0.19381 | 0.058045 | -0.16112 | -0.2289 | -0.091156 | -0.1909 | -0.082814 | -0.064425 | a5 | 2.0464 | -0.26848 | -0.038237 | -0.061186 | 0.080057 | -0.062896 | -0.046328 | 0.047811 | 0.047542 | -0.10822 | -0.025806 | -0.092434 | -0.067994 | 0.023314 | 0.15979 | -0.028912 | -0.0087675 | -0.013145 | -0.064479 | -0.048267 | 0.039762
a6 | 0.39563 | 0.055054 | -0.013801 | -0.025993 | 0.023477 | -0.0071092 | 0.025983 | 0.061751 | -0.077091 | -0.094542 | -0.039825 | -0.042274 | -0.066133 | -0.02823 | -0.006404 | -0.029465 | -0.014442 | -0.075952 | -0.0049251 | -0.29254 | -0.014908 | a6 | -0.061993 | -0.020899 | -0.052406 | 0.098881 | -0.0018421 | 0.089479 | -0.043302 | 0.02463 | 0.085883 | 0.091105 | 0.027939 | 0.052138 | -0.0069503 | 0.050253 | -0.0091117 | 0.045701 | 0.024152 | -0.032426 | 0.0097303 | -0.15684 | 0.015358
a7 | 1.3603 | -0.041012 | -0.037059 | -0.000299 | -0.013264 | 0.047369 | 0.064089 | 0.049624 | 0.037028 | -0.0091815 | -0.059179 | -0.053027 | 0.0046108 | -0.018821 | 0.072475 | 0.011783 | -0.020688 | 0.03555 | -0.054973 | 0.16916 | 0.051487 | a7 | 1.5727 | -0.068245 | -0.017133 | -0.018997 | 0.037345 | 0.060246 | 0.049725 | 0.040226 | 0.14756 | 0.051664 | 0.029881 | -0.0026404 | -0.0055311 | 0.016548 | 0.1215 | 0.023817 | 0.022179 | 0.035575 | -0.060259 | 0.051442 | 0.078492
a8 | 0.0076448 | 0.012557 | -0.01158 | -0.017008 | 0.016752 | 0.023972 | 0.012853 | 0.053586 | -0.0044769 | -0.02257 | -0.0049429 | 0.014389 | -0.0095054 | 0.02396 | 0.011514 | -0.0054928 | 0.017834 | -0.018181 | -0.0076205 | -0.010571 | 0.0043815 | a8 | -0.015878 | 0.0296 | 0.012787 | 0.092354 | 0.028707 | 0.063166 | 0.00034165 | 0.0053198 | 0.027737 | 0.027423 | 0.043489 | 0.013216 | 0.017562 | 0.053792 | 0.0321 | 0.040429 | 0.043757 | -0.023519 | -0.0055277 | -0.017324 | 0.025006
a9 | -0.32776 | -0.081106 | -0.050597 | -0.050368 | -0.039677 | -0.0089274 | -0.0042191 | 0.0075199 | -0.02248 | -0.032918 | 0.011581 | -0.036658 | -0.028357 | -0.027822 | 0.0036364 | -0.012509 | -0.029183 | -0.027822 | -0.036735 | -0.069645 | -0.03738 | a9 | -0.2666 | -0.039023 | -0.01652 | -0.031135 | -0.036643 | 0.0145 | -0.0069591 | 0.0087856 | -0.018904 | -0.016447 | 0.016321 | -0.024476 | -0.013837 | -0.037993 | 0.0060744 | -0.012309 | -0.00028437 | -0.015414 | -0.0367 | -0.041126 | -0.023852
a10 | 0.051749 | 0.076593 | 0.045855 | -0.00044719 | 0.035256 | 0.013523 | -0.0078405 | 0.018582 | -0.011873 | -0.062663 | -0.0099946 | -0.01936 | -0.018478 | 0.0064355 | -0.0034752 | -0.0007086 | 0.013073 | -0.018115 | 0.044153 | 0.018619 | 0.0075436 | a10 | -0.035499 | 0.0011601 | -0.0080573 | 0.066427 | -0.022137 | 0.02566 | -0.033617 | 0.02611 | -0.01054 | -0.006493 | 0.050199 | -0.0038367 | -0.0068444 | 0.034238 | 0.03117 | 0.052431 | 0.057177 | -0.05359 | -0.02384 | -0.046036 | 0.0064474
a11 | 0.13482 | -0.063561 | -0.065303 | -0.057247 | -0.048854 | -0.040926 | -0.040105 | -0.042486 | -0.041536 | -0.047936 | -0.045457 | -0.054911 | -0.046178 | -0.041782 | -0.033727 | -0.048218 | -0.05166 | -0.046312 | -0.057186 | -0.024144 | -0.041639 | a11 | 0.13708 | -0.066396 | -0.053563 | -0.047131 | -0.039965 | -0.033856 | -0.04081 | -0.039058 | -0.028334 | -0.03795 | -0.036665 | -0.045789 | -0.045235 | -0.044002 | -0.03245 | -0.041289 | -0.042216 | -0.047853 | -0.056995 | -0.033023 | -0.036969
a12 | -0.010006 | 0.0087107 | 0.001254 | -0.0043509 | 0.0033887 | -0.0015421 | -0.0010289 | 0.00090475 | -0.0071093 | -0.011804 | -0.0032323 | -0.0031488 | -0.0081317 | -0.003855 | -0.0045524 | -0.0043189 | -0.003162 | -0.0097388 | 0.00084307 | -0.01095 | -0.0033257 | a12 | -0.0050551 | -0.0019499 | -0.0023327 | 0.0095352 | -0.0034138 | 0.0063355 | -0.0071689 | 0.0018115 | -0.00096008 | 0.0029091 | 0.0060441 | -0.00015281 | -0.003149 | 0.0041533 | 0.00013366 | 0.0061318 | 0.0059066 | -0.0088014 | -0.0019131 | -0.0039296 | 0.00038497
a13 | -0.023261 | -0.013971 | -0.0083952 | -0.0098542 | -0.0058263 | -0.0061481 | -0.0037788 | -0.0055518 | -0.0061909 | -0.0058886 | -0.0033123 | -0.0081124 | -0.0079374 | -0.0068017 | -0.0034483 | -0.0055854 | -0.0085047 | -0.0068895 | -0.008748 | -0.0025784 | -0.0073452 | a13 | -0.034786 | -0.0094755 | -0.0049248 | -0.0094703 | -0.0060521 | -0.0022242 | -0.005812 | -0.0059621 | -0.0057747 | -0.0044744 | -0.0053168 | -0.0054259 | -0.0040595 | -0.0042799 | -0.0032507 | -0.0034959 | -0.0073414 | -0.0035822 | -0.0079171 | -0.0046385 | -0.0064542
a14 | 0.0097799 | -0.018438 | -0.018986 | -0.024228 | -0.020012 | -0.02091 | -0.02143 | -0.020762 | -0.023278 | -0.027118 | -0.02256 | -0.025233 | -0.025453 | -0.022146 | -0.023949 | -0.022867 | -0.021715 | -0.023189 | -0.017548 | -0.00045497 | -0.020448 | a14 | 0.0033077 | -0.021955 | -0.023354 | -0.01693 | -0.026289 | -0.018163 | -0.027377 | -0.021887 | -0.023693 | -0.021914 | -0.019043 | -0.020378 | -0.021323 | -0.016626 | -0.019672 | -0.018818 | -0.019974 | -0.02638 | -0.023846 | -0.0057246 | -0.020192
a15 | 0.0086944 | -0.012886 | -0.018288 | -0.013299 | -0.018267 | -0.010556 | -0.021836 | -0.0056397 | -0.018863 | -0.026744 | -0.022627 | -0.019928 | -0.0049733 | -0.018688 | -0.016012 | -0.019965 | -0.012324 | -0.020456 | -0.015747 | 0.011409 | -0.014771 | a15 | -0.0078717 | -0.011547 | -0.020163 | -0.001782 | -0.0084071 | -0.020917 | -0.0068684 | -0.007355 | -0.0065404 | -0.020329 | -0.00093479 | -0.021237 | -0.011109 | -0.020845 | -0.010673 | -0.019171 | -0.0035188 | -0.02302 | -0.015735 | -0.01741 | -0.012774
a16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | a16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
a17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | a17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Dx (mm) | 0.20295 | 0.063757 | -0.093946 | 0.053168 | -0.068905 | 0.099102 | -0.078562 | 0.21133 | -0.035973 | -0.14717 | -0.12504 | -0.056398 | 0.22351 | -0.05257 | -0.0045448 | -0.062709 | 0.073059 | -0.066303 | -0.018327 | 0.45682 |  | Dx (mm) | -0.087266 | 0.088043 | -0.09334 | 0.24843 | 0.06909 | -0.16614 | 0.15403 | 0.11991 | 0.11963 | -0.12716 | 0.21878 | -0.15195 | 0.05995 | -0.19094 | 0.0089414 | -0.16082 | 0.17868 | -0.1448 | -0.01192 | -0.094761
Dy (mm) | 0.10091 | 0.18073 | 0.09713 | 0.003097 | 0.076337 | 0.036556 | -0.022262 | 0.055327 | -0.029526 | -0.15326 | -0.030752 | -0.048605 | -0.028412 | 0.011859 | -0.0081151 | -0.0056593 | 0.034994 | -0.045727 | 0.099888 | 0.072876 |  | Dy (mm) | -0.08221 | 0.010607 | -0.02542 | 0.16831 | -0.046977 | 0.051219 | -0.068446 | 0.066553 | -0.019793 | -0.021472 | 0.12589 | -0.018344 | -0.012091 | 0.067635 | 0.070585 | 0.11008 | 0.1392 | -0.13219 | -0.055978 | -0.14153
Mag. len. (m) | 14.4427 | [20 empty] | Mag. len. (m) | 14.4468
TF (mT/KA) | 594.6013 | [20 empty] | TF (mT/KA) | 594.6209
Coil pos. (mm) | 0.016382 | [20 empty] | Coil pos. (mm) | 0.015878
Current (A) | 8.5 | [20 empty] | Current (A) | 8.5

## Harmonics
Normal:
Skew:
*10

## Work sheet
| Measured harmonics- Aperture 1 | [7 empty] | Measured harmonics- Aperture 2 | [8 empty] | Zero vector
| normal multipoles |  |  |  | skew multipoles |  |  |  | normal multipoles |  |  |  | skew multipoles
| all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19) |  | all positions (1:20) |  | central positions (2:19)
| average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma | average | sigma
1 | [17 empty] | 0
2 | -0.31165 | 1.0617 | -0.18868 | 0.65127 | -1.1339 | 3.2777 | -1.2617 | 0.49228 | 0.20407 | 1.168 | 0.33589 | 0.69578 | 0.84994 | 1.7196 | 0.8117 | 0.45745 |  | 0
3 | -4.3362 | 5.195 | -4.8054 | 0.28547 | -0.50765 | 0.75328 | -0.4158 | 0.29084 | -3.8712 | 5.2089 | -4.379 | 0.30458 | -0.23914 | 0.45814 | -0.20662 | 0.30515 |  | 0
4 | -0.18073 | 0.50904 | -0.081453 | 0.096751 | -0.076091 | 0.64304 | -0.13533 | 0.13818 | -0.033194 | 0.22493 | -0.015583 | 0.11031 | 0.045413 | 0.32621 | -0.0078974 | 0.17264 |  | 0
5 | -0.20974 | 2.0115 | 0.2097 | 0.12863 | -0.064425 | 0.41385 | -0.1248 | 0.13898 | -0.019473 | 2.0291 | 0.40562 | 0.14123 | 0.039762 | 0.47202 | -0.029454 | 0.089263 |  | 0
6 | -0.011427 | 0.063461 | -0.0094777 | 0.05353 | -0.014908 | 0.12099 | -0.023819 | 0.029531 | -0.031906 | 0.090047 | -0.027064 | 0.040388 | 0.015358 | 0.062934 | 0.01647 | 0.042593 |  | 0
7 | 0.91387 | 0.23705 | 0.9191 | 0.041 | 0.051487 | 0.30695 | 9.3614e-06 | 0.042115 | 0.95392 | 0.2504 | 0.95569 | 0.044144 | 0.078492 | 0.34936 | 0.024909 | 0.054522 |  | 0
8 | -0.0020741 | 0.044123 | -0.003486 | 0.018564 | 0.0043815 | 0.018627 | 0.0013829 | 0.021505 | -0.003887 | 0.036203 | -0.0067207 | 0.011521 | 0.025006 | 0.028922 | 0.021564 | 0.016636 |  | 0
9 | 0.55921 | 0.081136 | 0.57496 | 0.0055487 | -0.03738 | 0.071134 | -0.02693 | 0.023381 | 0.56028 | 0.094631 | 0.57678 | 0.0052399 | -0.023852 | 0.058911 | -0.016259 | 0.018482 |  | 0
10 | 0.0057185 | 0.058137 | -6.2387e-10 | 7.3879e-09 | 0.0075436 | 0.031416 | 2.3013e-09 | 5.8916e-09 | 0.0035473 | 0.062285 | -2.3405e-10 | 1.3234e-08 | 0.0064474 | 0.036065 | 9.8683e-10 | 4.8082e-09 |  | 0
11 | 0.7307 | 0.045487 | 0.74006 | 0.0038183 | -0.041639 | 0.041839 | -0.048732 | 0.0084919 | 0.73044 | 0.049328 | 0.7403 | 0.005938 | -0.036969 | 0.041197 | -0.043579 | 0.0093449 |  | 0
12 | -0.00057576 | 0.0072812 | -0.00098817 | 0.0016586 | -0.0033257 | 0.0051826 | -0.0035508 | 0.0020401 | 0.001899 | 0.0091159 | 0.001007 | 0.0021306 | 0.00038497 | 0.0049921 | -0.00048206 | 0.0019688 |  | 0
13 | 0.080034 | 0.0096035 | 0.081251 | 0.0012612 | -0.0073452 | 0.0045249 | -0.0070394 | 0.0026272 | 0.079312 | 0.008274 | 0.080074 | 0.0012146 | -0.0064542 | 0.006838 | -0.005292 | 0.0020775 |  | 0
14 | -0.010073 | 0.0045817 | -0.010372 | 0.0010244 | -0.020448 | 0.0087712 | -0.023031 | 0.00094925 | -0.0066579 | 0.006055 | -0.0075889 | 0.0010964 | -0.020192 | 0.0070933 | -0.022899 | 0.0011592 |  | 0
15 | 0.044413 | 0.010143 | 0.046432 | 0.0036631 | -0.014771 | 0.0098085 | -0.016506 | 0.0057124 | 0.044611 | 0.010298 | 0.046441 | 0.004004 | -0.012774 | 0.0072274 | -0.012786 | 0.0074634 |  | 0
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0
Specified harmonics
| normal harmonics |  |  |  |  | skew harmonics
| average | uncertain | sigma | min | max | average | uncertain | sigma | min | max |  |  | Coil positioning estimate
1 | 0 | 10 | 5 | -15 | 15 | 0.75 | 5 | 0 | 0.75 | 0.75 |  |  | n | n^2 | Ap. 1 | Ap. 2
2 | -1.4 | 0.85 | 0.68 | -3.44 | 0.64 | 0 | 0.51 | 1.7 | -5.1 | 5.1 |  |  | 2 | 4 | -0.56877 | -0.57241
3 | -4 | 0.87 | 1.45 | -8.35 | 0.35 | 0.08 | 0.87 | 0.43 | -1.21 | 1.37 |  |  | 3 | 9 | -1.2443 | -1.1879
4 | 0.22 | 0.34 | 0.49 | -1.25 | 1.69 | 0 | 0.13 | 0.49 | -1.47 | 1.47 |  |  | 4 | 16 | -2.1574 | -1.9805
5 | 0 | 0.42 | 0.42 | -1.26 | 1.26 | 0.01 | 0.42 | 0.33 | -0.98 | 1 |  |  | 5 | 25 | -2.0121 | -2.1868
6 | -0.01 | 0.06 | 0.09 | -0.28 | 0.26 | 0 | 0.06 | 0.14 | -0.42 | 0.42 |  |  | 6 | 36 | -3.2249 | -3.1826
7 | 0.32 | 0 | 0.22 | -0.34 | 0.98 | 0.02 | 0 | 0.07 | -0.19 | 0.23 |  |  | 7 | 49 | -3.1808 | -3.0147
8 | 0 | 0 | 0.04 | -0.12 | 0.12 | 0 | 0 | 0.08 | -0.24 | 0.24 |  |  | 8 | 64 | -3.913 | -4.2799
9 | 0.13 | 0 | 0.07 | -0.08 | 0.34 | -0.01 | 0 | 0.07 | -0.22 | 0.2 |  |  | 9 | 81 | -4.475 | -4.6212
10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  |  | sigma (mm) |  | 0.016382 | 0.015878
11 | 0.53 | 0 | 0 | 0.53 | 0.53 | 0 | 0 | 0 | 0 | 0
12 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
13 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
14 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
15 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| Random harmonics | [6 empty] | Expected harmonics (no uncertainty)
| d (mm) |  |  |  | sig(n)=d alpha beta^n |  |  | normal |  |  | skew
| 0.1 | 0.025 | 0.006 |  |  |  |  | min | ave | max | min | ave | max
1 | 4.6023 | 1.1506 | 0.27614 |  | Scaling law constants |  |  | -15 | 0 | 15 | 0.75 | 0.75 | 0.75
2 | 2.8314 | 0.70784 | 0.16988 |  | alpha | 73.9 |  | -3.44 | -1.4 | 0.64 | -5.1 | 0 | 5.1
3 | 1.7207 | 0.43017 | 0.10324 |  | beta | 0.6266 |  | -8.35 | -4 | 0.35 | -1.21 | 0.08 | 1.37
4 | 1.033 | 0.25824 | 0.061978 |  | gamma | 0.9939 |  | -1.25 | 0.22 | 1.69 | -1.47 | 0 | 1.47
5 | 0.61258 | 0.15315 | 0.036755 |  |  |  |  | -1.26 | 0 | 1.26 | -0.98 | 0.01 | 1
6 | 0.35886 | 0.089715 | 0.021532 |  |  |  |  | -0.28 | -0.01 | 0.26 | -0.42 | 0 | 0.42
7 | 0.20767 | 0.051917 | 0.01246 |  |  |  |  | -0.34 | 0.32 | 0.98 | -0.19 | 0.02 | 0.23
8 | 0.11871 | 0.029678 | 0.0071228 |  |  |  |  | -0.12 | 0 | 0.12 | -0.24 | 0 | 0.24
9 | 0.067037 | 0.016759 | 0.0040222 |  |  |  |  | -0.08 | 0.13 | 0.34 | -0.22 | -0.01 | 0.2
10 | 0.037395 | 0.0093488 | 0.0022437 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
11 | 0.020607 | 0.0051516 | 0.0012364 |  |  |  |  | 0.53 | 0.53 | 0.53 | 0 | 0 | 0
12 | 0.011217 | 0.0028042 | 0.00067302 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
13 | 0.0060316 | 0.0015079 | 0.0003619 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
14 | 0.0032039 | 0.00080097 | 0.00019223 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
15 | 0.0016811 | 0.00042029 | 0.00010087 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
16 | 0.0008714 | 0.00021785 | 5.2284e-05 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
17 | 0.00044619 | 0.00011155 | 2.6771e-05 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | ten times !
Measured harmonics along the axis with local feed down - Aperture 1 - Normal
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20
2 | -4.531 | -0.4643 | -1.5185 | -1.1478 | 0.099631 | 0.30892 | 0.60789 | 0.90849 | -0.020578 | -0.63032 | -0.18936 | 0.68569 | -0.26989 | -1.1121 | -0.59177 | -0.019069 | 0.29182 | -0.2274 | -0.10755 | 0.55227 | -0.31165
3 | 16.8192 | -4.5191 | -4.9813 | -4.899 | -4.9392 | -4.4664 | -4.9833 | -4.9106 | -5.0299 | -4.8977 | -4.2314 | -4.3962 | -4.4592 | -4.7142 | -5.0862 | -5.1363 | -4.9359 | -5.233 | -4.679 | -12.382 | -4.3362
4 | -1.7418 | -0.11502 | 0.020571 | -0.031186 | -0.043365 | 0.035309 | -0.016739 | -0.10933 | -0.10944 | -0.023988 | 0.027036 | -0.15903 | -0.12689 | -0.14893 | -0.15649 | 0.038245 | -0.22052 | -0.3113 | -0.015094 | -0.55483 | -0.18073
5 | -7.3713 | 0.023877 | 0.10417 | 0.16975 | 0.11116 | 0.24494 | 0.45869 | 0.42567 | 0.3444 | 0.27118 | 0.29709 | 0.26745 | 0.32841 | 0.13708 | 0.21457 | 0.12063 | 0.11699 | 0.1247 | 0.013815 | -5.002 | -0.20974
6 | -0.11852 | -0.030516 | -0.070006 | -0.037425 | -0.047654 | 0.037243 | 0.090421 | 0.084002 | 0.036273 | -0.0086724 | -0.071189 | -0.052849 | -0.081778 | -0.02713 | -0.035395 | 0.023744 | -0.054028 | 0.033315 | 0.041047 | 0.0059067 | -0.011427
7 | 1.5356 | 0.79172 | 0.90754 | 0.92311 | 0.92512 | 0.90624 | 0.97123 | 0.95611 | 0.95795 | 0.90965 | 0.95188 | 0.94539 | 0.96225 | 0.93086 | 0.89565 | 0.90973 | 0.9019 | 0.91848 | 0.87892 | 0.077465 | 0.91387
8 | -0.11243 | 0.012407 | 0.0034697 | 0.025591 | 0.026795 | -0.0058419 | -0.011481 | -0.029753 | 0.0069327 | 0.004106 | -0.011262 | -0.032986 | -0.026835 | -0.00063139 | 0.0041089 | 0.0028365 | -0.019954 | -0.02761 | 0.01736 | 0.050282 | -0.0020741
9 | 0.27671 | 0.57091 | 0.56518 | 0.56806 | 0.57419 | 0.58555 | 0.58209 | 0.57679 | 0.56974 | 0.56823 | 0.5806 | 0.5808 | 0.57938 | 0.57307 | 0.57771 | 0.57605 | 0.57515 | 0.5699 | 0.57594 | 0.31452 | 0.55921
10 | -1.4582e-08 | 5.1459e-09 | -6.1958e-09 | -2.6632e-09 | 3.8718e-09 | 3.2991e-09 | -3.9208e-09 | -9.1558e-09 | -5.5164e-09 | -1.9764e-08 | -1.8233e-09 | 3.1003e-09 | 1.6105e-08 | 4.177e-09 | -2.1054e-10 | 2.0954e-09 | 3.9264e-09 | -1.438e-09 | -2.263e-09 | -1.1735e-08 | 0.0057185
11 | 0.59948 | 0.74356 | 0.73907 | 0.73727 | 0.74105 | 0.73946 | 0.73993 | 0.73181 | 0.73428 | 0.74104 | 0.73671 | 0.74082 | 0.74061 | 0.73718 | 0.74691 | 0.74691 | 0.74279 | 0.74066 | 0.74101 | 0.58147 | 0.7307
12 | -0.07942 | 0.029967 | -0.019309 | -0.018873 | 0.0037717 | -0.0096034 | -0.021814 | -0.013734 | -0.012263 | -0.042781 | -0.012232 | -0.00099782 | 0.0067709 | 0.0036813 | -0.01406 | -0.025162 | -0.026447 | -0.014723 | 0.0099372 | -0.0057691 | -0.0057576 | ten times!
13 | 0.73531 | 0.79977 | 0.81052 | 0.82827 | 0.8275 | 0.80814 | 0.79723 | 0.79591 | 0.80662 | 0.81517 | 0.81651 | 0.8124 | 0.78812 | 0.81295 | 0.83393 | 0.82605 | 0.81125 | 0.80759 | 0.82721 | 0.40791 | 0.80034 | ten times!
14 | -0.064802 | -0.12496 | -0.11504 | -0.11749 | -0.10739 | -0.09985 | -0.10551 | -0.10047 | -0.10113 | -0.08625 | -0.091609 | -0.097512 | -0.11141 | -0.098628 | -0.10968 | -0.11118 | -0.10515 | -0.089243 | -0.094368 | -0.087318 | -0.10073 | ten times!
15 | 0.12743 | 0.40838 | 0.40229 | 0.48208 | 0.4198 | 0.46987 | 0.47569 | 0.46971 | 0.48518 | 0.52575 | 0.46305 | 0.48048 | 0.50394 | 0.46199 | 0.50498 | 0.4672 | 0.46928 | 0.47756 | 0.39057 | 0.18475 | 0.44413 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Measured harmonics along the axis with local feed down- Aperture 1 - Skew
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1
2 | 10.3377 | -1.1596 | -1.2846 | -1.2253 | -1.0792 | -1.0739 | -0.58615 | -0.45234 | -1.4055 | -0.9941 | -1.1443 | -2.6836 | -1.7127 | -1.4094 | -1.5464 | -1.5304 | -1.6422 | -0.88835 | -0.89179 | -9.1205 | -1.1339
3 | -3.5682 | -1.2388 | -0.1956 | -0.76932 | -0.42449 | -0.53771 | -0.086363 | -0.15057 | -0.33645 | -0.67311 | -0.16974 | -0.2132 | -0.11724 | -0.65392 | -0.26079 | -0.34244 | -0.46245 | -0.28167 | -0.57064 | 0.00076326 | -0.50765
4 | 2.5908 | 0.044277 | -0.12257 | -0.12919 | -0.10323 | -0.16165 | -0.052541 | 0.12007 | -0.011728 | -0.46333 | -0.35386 | -0.22021 | -0.24879 | -0.22563 | -0.099382 | -0.12805 | -0.13622 | -0.143 | -0.00084107 | -0.91605 | -0.076091
5 | 1.6096 | -0.43371 | -0.23161 | -0.24716 | -0.072582 | -0.087625 | 0.037448 | 0.13718 | 0.082306 | -0.16102 | -0.098562 | -0.19764 | -0.16293 | -0.1941 | 0.057949 | -0.16166 | -0.22827 | -0.092471 | -0.19203 | -0.043894 | -0.064425
6 | 0.24372 | 0.0054752 | -0.046182 | -0.026997 | -0.0017486 | -0.020465 | 0.0354 | 0.039355 | -0.066638 | -0.045769 | -0.032081 | -0.027088 | -0.056816 | -0.032476 | -0.0037223 | -0.027387 | -0.025044 | -0.060282 | -0.036274 | -0.32244 | -0.014908
7 | 1.3626 | -0.044039 | -0.03674 | 2.248e-05 | -0.013091 | 0.046063 | 0.064201 | 0.044314 | 0.036925 | -0.012925 | -0.060009 | -0.053678 | 0.005894 | -0.01823 | 0.072506 | 0.011608 | -0.021229 | 0.034188 | -0.055612 | 0.1673 | 0.051487
8 | 0.025614 | -0.03304 | -0.039627 | -0.016577 | -0.0051462 | 0.01428 | 0.018809 | 0.037565 | 0.0030607 | 0.01615 | 0.0041912 | 0.026707 | 0.001373 | 0.020071 | 0.013728 | -0.0043254 | 0.0093491 | -0.0067728 | -0.034902 | -0.0076747 | 0.0043815
9 | -0.33672 | -0.087004 | -0.046407 | -0.050393 | -0.037265 | -0.010489 | -0.0049545 | 0.0027304 | -0.02296 | -0.043356 | 0.010023 | -0.037909 | -0.024714 | -0.027502 | 0.0036188 | -0.012615 | -0.0303 | -0.029187 | -0.036062 | -0.08026 | -0.03738
10 | 4.7886e-09 | 1.6611e-08 | 6.1262e-09 | -7.8873e-11 | 8.4738e-10 | -1.529e-09 | 1.9014e-09 | 1.7674e-09 | 2.6427e-09 | 1.6875e-08 | -9.9603e-10 | 2.7648e-09 | -5.7741e-09 | -1.7529e-09 | -1.0806e-10 | -2.9911e-10 | -2.9693e-09 | 1.7263e-09 | 3.6681e-09 | 1.4158e-09 | 0.0075436
11 | 0.13587 | -0.064912 | -0.064794 | -0.057099 | -0.048534 | -0.040942 | -0.040254 | -0.042996 | -0.041767 | -0.050376 | -0.045889 | -0.055139 | -0.044766 | -0.041893 | -0.033749 | -0.048416 | -0.05155 | -0.046892 | -0.05721 | -0.021494 | -0.041639 | ten times!
12 | -0.11801 | -0.01136 | -0.050212 | -0.041614 | -0.014506 | -0.032011 | 0.00021812 | -0.013996 | -0.055895 | -0.037175 | -0.017336 | -0.0069777 | -0.053166 | -0.047935 | -0.040863 | -0.042339 | -0.047316 | -0.074605 | -0.052058 | -0.12174 | -0.033257 | ten times!
13 | -0.24387 | -0.11907 | -0.087822 | -0.088464 | -0.061818 | -0.044005 | -0.05305 | -0.021283 | -0.071024 | -0.11094 | -0.058104 | -0.096757 | -0.0363 | -0.075861 | -0.036014 | -0.06741 | -0.070952 | -0.084954 | -0.08327 | -0.02282 | -0.073452 | ten times!
14 | 0.072678 | -0.2384 | -0.23619 | -0.23768 | -0.23688 | -0.21463 | -0.2197 | -0.2192 | -0.22657 | -0.23724 | -0.23717 | -0.24236 | -0.23359 | -0.23406 | -0.23671 | -0.2368 | -0.22325 | -0.22508 | -0.20998 | -0.058558 | -0.20448 | ten times!
15 | 0.086944 | -0.12886 | -0.18288 | -0.13299 | -0.18267 | -0.10556 | -0.21836 | -0.056397 | -0.18863 | -0.26744 | -0.22627 | -0.19928 | -0.049733 | -0.18688 | -0.16012 | -0.19965 | -0.12324 | -0.20456 | -0.15747 | 0.11409 | -0.14771 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Measured harmonics along the axis with local feed down - Aperture 2 - Normal
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1
2 | -3.7675 | 1.0879 | 1.2466 | 0.3612 | 1.1269 | 0.7278 | 1.4499 | 1.2232 | -0.14444 | -0.55593 | -0.065881 | 0.2282 | 0.29892 | -0.55306 | 0.74061 | 0.16861 | -0.36867 | -0.65234 | -0.27341 | 0.7778 | 0.20407
3 | 17.4008 | -4.5968 | -4.6185 | -4.1129 | -4.6887 | -4.8381 | -4.7604 | -4.6368 | -4.3273 | -4.2618 | -4.3144 | -4.1684 | -3.8149 | -4.101 | -4.3157 | -3.9457 | -4.0581 | -4.6502 | -4.6117 | -11.3931 | -3.8712
4 | -1.0269 | 0.011025 | 0.23487 | 0.12459 | -0.045883 | -0.04456 | -0.044174 | -0.1892 | -0.13183 | -0.068597 | 0.030825 | -0.068556 | -0.015946 | -0.091347 | -0.19361 | 0.11317 | -0.02878 | 0.025093 | 0.10242 | 0.16583 | -0.033194
5 | -7.1877 | 0.24298 | 0.063657 | 0.25375 | 0.35471 | 0.57393 | 0.58082 | 0.482 | 0.46665 | 0.51464 | 0.51383 | 0.43804 | 0.46501 | 0.47607 | 0.52602 | 0.46682 | 0.34154 | 0.33892 | 0.2018 | -4.9413 | -0.019473
6 | -0.28981 | -0.018975 | 0.061902 | 0.0046855 | -0.009459 | -0.04244 | -0.058359 | -0.036239 | -0.0073777 | -0.028938 | 0.00090382 | -0.045018 | -0.053725 | -0.12576 | -0.017992 | -0.009007 | -0.073402 | -0.047794 | 0.019841 | -0.0027467 | -0.031906
7 | 1.6544 | 0.86617 | 0.90699 | 0.9362 | 0.95116 | 0.90014 | 1.0187 | 1.0061 | 1.0033 | 0.95402 | 0.9364 | 1.0045 | 0.98283 | 1.0135 | 0.91027 | 0.97975 | 0.92588 | 0.96019 | 0.9463 | 0.13704 | 0.95392
8 | 0.018519 | -0.038206 | -0.0086643 | -0.010841 | 0.0055584 | -0.0053427 | -0.012549 | -0.019404 | -0.0067956 | 0.0062502 | -0.0072084 | -0.010595 | -0.020136 | -0.0060011 | -0.0027198 | 0.005006 | -0.0012401 | 0.0022377 | 0.0096789 | -0.0067624 | -0.003887
9 | 0.21893 | 0.57252 | 0.56587 | 0.57669 | 0.58423 | 0.56945 | 0.58166 | 0.57705 | 0.57785 | 0.57153 | 0.58073 | 0.58141 | 0.58305 | 0.57174 | 0.57276 | 0.58059 | 0.58263 | 0.57626 | 0.57605 | 0.34347 | 0.56028
10 | -4.6507e-09 | -2.7657e-09 | 3.1516e-09 | 3.4526e-08 | 1.6645e-09 | 1.7037e-08 | 1.9563e-09 | -1.4779e-08 | 1.2594e-09 | -1.3366e-08 | -7.4924e-09 | -7.3426e-09 | 2.7163e-09 | -2.2543e-08 | 4.6959e-10 | -1.523e-08 | 1.3679e-08 | 4.8902e-09 | -2.0439e-09 | 9.8111e-10 | 0.0035473
11 | 0.58946 | 0.73683 | 0.73515 | 0.73895 | 0.74213 | 0.74093 | 0.74228 | 0.7366 | 0.73358 | 0.73827 | 0.74024 | 0.73424 | 0.73789 | 0.73565 | 0.7548 | 0.74885 | 0.75138 | 0.7414 | 0.73616 | 0.5751 | 0.73044
12 | 0.17665 | -0.010632 | -0.016461 | 0.0046985 | 0.02897 | -0.0097574 | 0.034202 | 0.029706 | -0.0074778 | 0.0098579 | -0.00022154 | 0.024961 | 0.015621 | 0.037201 | 0.049439 | 0.00079022 | 0.025835 | -0.016284 | -0.01918 | -0.021327 | 0.01899 | ten times!
13 | 0.70036 | 0.80179 | 0.795 | 0.81366 | 0.79456 | 0.78554 | 0.80057 | 0.77964 | 0.80595 | 0.79017 | 0.79602 | 0.80505 | 0.79731 | 0.79117 | 0.82302 | 0.80372 | 0.8252 | 0.79247 | 0.81247 | 0.45874 | 0.79312 | ten times!
14 | 0.06837 | -0.097233 | -0.072559 | -0.078141 | -0.062216 | -0.075721 | -0.092207 | -0.073334 | -0.073463 | -0.064256 | -0.058574 | -0.071544 | -0.072228 | -0.076285 | -0.061409 | -0.090067 | -0.079151 | -0.077036 | -0.090584 | -0.040942 | -0.066579 | ten times!
15 | 0.051604 | 0.49489 | 0.44595 | 0.42263 | 0.51758 | 0.43406 | 0.53205 | 0.4567 | 0.50213 | 0.4906 | 0.42237 | 0.44907 | 0.47927 | 0.40244 | 0.44454 | 0.414 | 0.43605 | 0.5182 | 0.49682 | 0.32659 | 0.44611 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Measured harmonics along the axis with local feed down - Aperture 2 - Skew
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
1
2 | 6.6734 | 1.137 | 1.0239 | 1.3743 | 0.89285 | 1.8939 | 1.3409 | 0.58987 | 0.67264 | 0.97851 | 0.62741 | 0.62551 | 0.85916 | 0.77566 | 0.75002 | 0.55146 | 0.15571 | -0.10724 | 0.46901 | -3.864 | 0.84994
3 | -1.6709 | -0.38405 | 0.067219 | -0.070091 | 0.37666 | -0.0024378 | -0.4267 | -0.55487 | -0.074584 | -0.88684 | 0.090054 | -0.2116 | -0.18893 | 0.11548 | -0.0080834 | -0.26455 | -0.29261 | -0.43154 | -0.57172 | 0.33043 | -0.23914
4 | 1.1412 | -0.13437 | -0.35838 | 0.0039812 | 0.034259 | 0.14915 | -0.10692 | -0.081304 | 0.22108 | 0.088836 | 0.12426 | 0.23216 | 0.19299 | 0.10556 | 0.012083 | -0.16844 | -0.21909 | -0.26127 | 0.023277 | 0.10857 | 0.045413
5 | 2.0377 | -0.26796 | -0.039438 | -0.072761 | 0.080293 | -0.057071 | -0.0444 | 0.046837 | 0.044731 | -0.10526 | -0.030272 | -0.090638 | -0.067989 | 0.030008 | 0.16019 | -0.024627 | -0.0093803 | -0.01823 | -0.064203 | -0.052836 | 0.039762
6 | 0.034436 | -0.022019 | -0.04482 | 0.04495 | 0.013038 | 0.076701 | -0.021349 | -0.00071643 | 0.086673 | 0.10067 | -0.016085 | 0.058522 | -0.0026393 | 0.027104 | -0.032149 | 0.0089584 | -0.022786 | 0.014237 | 0.02817 | -0.14828 | 0.015358
7 | 1.5721 | -0.069256 | -0.016995 | -0.032609 | 0.036988 | 0.065636 | 0.05054 | 0.039607 | 0.1464 | 0.052849 | 0.023245 | -0.002212 | -0.0059833 | 0.022359 | 0.1214 | 0.02824 | 0.016204 | 0.0321 | -0.06016 | 0.049324 | 0.078492
8 | -0.018374 | 0.02835 | 0.018853 | 0.049612 | 0.042825 | 0.05044 | 0.019715 | -0.013326 | 0.034212 | 0.032251 | 0.0069296 | 0.01653 | 0.021274 | 0.03196 | 0.013087 | 0.0092775 | 0.0053794 | 0.011365 | 0.0094237 | 0.0036036 | 0.025006
9 | -0.26926 | -0.039293 | -0.017473 | -0.04993 | -0.03511 | 0.018695 | -0.0018376 | 0.005242 | -0.017699 | -0.017518 | 0.0039647 | -0.025427 | -0.013455 | -0.031634 | 0.0057281 | -0.0038691 | -0.011776 | -0.024204 | -0.037059 | -0.046043 | -0.023852
10 | 3.4102e-09 | -7.8852e-10 | -1.0319e-09 | 8.7871e-09 | -5.6958e-10 | -3.2321e-09 | 4.145e-09 | -6.2904e-09 | -3.9929e-09 | 1.3353e-09 | 1.2005e-08 | 2.0827e-09 | -2.4172e-10 | 8.7124e-11 | -8.6992e-10 | -1.7469e-09 | 6.3261e-09 | 6.0024e-09 | -4.2441e-09 | -6.9846e-09 | 0.0064474
11 | 0.1374 | -0.066312 | -0.053819 | -0.050437 | -0.039614 | -0.032831 | -0.039579 | -0.039626 | -0.028184 | -0.037798 | -0.038586 | -0.04588 | -0.045074 | -0.043178 | -0.032708 | -0.040001 | -0.044157 | -0.049531 | -0.057115 | -0.033782 | -0.036969
12 | -0.031446 | -0.019596 | -0.012246 | 0.012462 | -0.0052151 | 0.031508 | -0.027545 | -0.013884 | 0.0064538 | 0.037173 | -0.0039898 | 0.0032365 | -0.023004 | -0.0033107 | -0.039976 | -0.0075546 | -0.014701 | -0.018522 | 0.01194 | 0.0034383 | 0.0038497 | ten times!
13 | -0.34192 | -0.079484 | -0.068076 | -0.063928 | -0.047926 | -0.040371 | -0.02737 | -0.038321 | -0.036452 | -0.06801 | -0.023088 | -0.079775 | -0.031279 | -0.060457 | -0.028359 | -0.047367 | -0.044956 | -0.081601 | -0.085742 | -0.059919 | -0.064542 | ten times!
14 | 0.030914 | -0.2155 | -0.2397 | -0.22424 | -0.23809 | -0.22856 | -0.23506 | -0.23664 | -0.2223 | -0.23175 | -0.23254 | -0.22358 | -0.20297 | -0.22145 | -0.22178 | -0.25111 | -0.24455 | -0.23484 | -0.2171 | -0.032768 | -0.20192 | ten times!
15 | -0.078717 | -0.11547 | -0.20163 | -0.01782 | -0.084071 | -0.20917 | -0.068684 | -0.07355 | -0.065404 | -0.20329 | -0.0093479 | -0.21237 | -0.11109 | -0.20845 | -0.10673 | -0.19171 | -0.035188 | -0.2302 | -0.15735 | -0.1741 | -0.12774 | ten times!
16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ten times!
Relative main field module along the axis
| position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
|  | 0.00057159 | -2.1984e-05 | -0.00021984 | -2.1984e-05 | -0.00021984 | -0.00021984 | -0.0004177 | -0.0004177 | -2.1984e-05 | -2.1984e-05 | -2.1984e-05 | -2.1984e-05 | -0.00021984 | 0.00017587 | 0.00017587 | 0.00017587 | 0.00017587 | 0.00057159 |  | 594.6013
|  | 0.0005386 | -5.4959e-05 | -5.4959e-05 | -0.00025281 | -5.4959e-05 | -0.00025281 | -0.00025281 | -0.00025281 | -0.00025281 | -0.00025281 | -0.00025281 | -0.00025281 | -5.4959e-05 | 0.00034075 | 0.00034075 | 0.00034075 | 0.00014289 | 0.0005386 |  | 594.6209
Check Computation for Aperture1
Magnetic Length
mbh dB/B | -3701.7067
| mean 3:18 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
dB/B | 594.5588 | -3586.9404 | 6.4309 | 0.49468 | -1.484 | 0.49468 | -1.484 | -1.484 | -3.4628 | -3.4628 | 0.49468 | 0.49468 | 0.49468 | 0.49468 | -1.484 | 2.4734 | 2.4734 | 2.4734 | 2.4734 | 6.4309 | -3816.4729
| mean 2:19 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
D Angle | -0.13488 | 4.9564 | -0.72388 | -0.5681 | -0.12748 | 0.45246 | 0.21846 | -0.28813 | -0.82785 | -0.49178 | 0.34571 | 0.70767 | 0.21312 | 0.57518 | 0.84237 | 0.6143 | -0.12449 | -0.17126 | -0.15971 | -0.48656 | -0.92702
db2 | -0.18374 | -4.0906 | -0.28808 | -1.2775 | -0.99445 | 0.3272 | 0.44293 | 0.83746 | 0.97097 | 0.18328 | -0.37348 | 0.056072 | 0.8974 | -0.20399 | -0.89828 | -0.40556 | 0.20234 | 0.43504 | -0.0043572 | 0.092985 | -0.11639
db3 | -4.8065 | 16.6984 | 0.28506 | -0.1732 | -0.09275 | -0.1308 | 0.3419 | -0.17672 | -0.1087 | -0.22279 | -0.10285 | 0.57278 | 0.41008 | 0.3416 | 0.094162 | -0.27969 | -0.33032 | -0.13139 | -0.42394 | 0.12756 | -12.4569
db4 | -0.079944 | -2.1333 | -0.016347 | 0.10349 | 0.051058 | 0.036103 | 0.12176 | 0.05494 | -0.0096595 | -0.031829 | 0.041061 | 0.097431 | -0.084882 | -0.030836 | -0.070153 | -0.076661 | 0.1162 | -0.13669 | -0.23427 | 0.069278 | -1.0901
db5 | 0.21023 | -7.3864 | -0.18947 | -0.10294 | -0.040774 | -0.098181 | 0.036773 | 0.24715 | 0.22403 | 0.13325 | 0.059135 | 0.090685 | 0.057817 | 0.11726 | -0.07237 | 0.0043681 | -0.089717 | -0.093747 | -0.086792 | -0.19647 | -4.9933
db6 | -0.010796 | -0.057414 | 0.00071175 | -0.088071 | -0.0092898 | -0.058979 | 0.079144 | 0.074758 | 0.16524 | 0.035293 | -0.045668 | -0.10314 | -0.061809 | 0.0052179 | -0.03353 | -0.025828 | 0.014421 | -0.019726 | 0.02317 | 0.048095 | 0.01799
db7 | 0.91984 | 1.5294 | -0.1281 | -0.011092 | 0.0041676 | 0.0044814 | -0.013076 | 0.052578 | 0.037519 | 0.038096 | -0.0093307 | 0.034317 | 0.026976 | 0.04551 | 0.01122 | -0.024162 | -0.009752 | -0.018186 | -0.00044763 | -0.040716 | 0.10193
db8 | -0.0041349 | -0.070051 | 0.040384 | -0.015047 | 0.044029 | 0.01375 | 0.02585 | -0.02896 | 0.032523 | 0.0011084 | -0.033505 | -0.041328 | -0.045112 | 0.039037 | -0.010533 | 0.0070221 | -0.010088 | 0.0044708 | -0.041875 | 0.018274 | 0.12837
db9 | 0.57624 | 0.2808 | -0.01149 | -0.011571 | -0.0075269 | -0.0024605 | 0.011356 | 0.0072002 | 0.010364 | -0.0063747 | -0.0077533 | 0.0078399 | 0.0048478 | 0.014318 | -0.0025852 | 0.0014597 | 0.00073474 | -5.2456e-05 | -0.0057181 | -0.0025877 | 0.35211
db10 | -0.0015919 | 0.063706 | 0.03621 | -0.035538 | 0.024756 | -0.02625 | 0.045599 | -0.033143 | 0.094127 | -0.014666 | -0.066891 | -0.053464 | -0.024556 | 0.098419 | -0.020917 | -0.00056586 | -0.026129 | 0.034584 | -0.028539 | -0.0030358 | 0.15841
db11 | 0.74031 | 0.60066 | 0.0025732 | -0.00090771 | -0.0029968 | 0.00072049 | -0.00050101 | -5.7506e-05 | -0.0070618 | -0.0060945 | 0.00081609 | -0.0029783 | 0.00053799 | 0.0020023 | -0.0030165 | 0.0065759 | 0.0068326 | 0.0026267 | 0.00029589 | 0.00063299 | 0.58567
db12 | -0.001187 | 0.004028 | 0.0096616 | -0.0056994 | 0.0024173 | -0.0022342 | 0.0060485 | -0.0055014 | 0.011996 | -0.0022116 | -0.012081 | -0.0073905 | -0.0024099 | 0.013999 | -0.0014295 | -0.00050591 | -0.0050325 | 0.002945 | -0.0042787 | 0.0017085 | 0.012439
db13 | 0.081697 | 0.071989 | 0.00041715 | 0.0017099 | 0.00079659 | 0.0028518 | -0.00074051 | -0.0014957 | -0.0014882 | -0.0012306 | -0.0012903 | 0.00081073 | -0.00084335 | -0.0037766 | 0.00026829 | 0.001587 | 0.0014627 | -0.00040028 | -0.0011252 | 0.0024865 | 0.039955
db14 | -0.010513 | -0.0050729 | 0.0020783 | -0.0026412 | 0.0009078 | -0.0014608 | 0.0046801 | -0.0035162 | 0.0088972 | -0.0014961 | -0.0078602 | -0.0039896 | -0.0022679 | 0.008531 | -0.0011678 | -0.00075173 | -0.0031113 | 0.0031764 | -0.0017897 | 0.0017817 | -0.0024662
db15 | 0.046432 | 0.012743 | -0.005594 | -0.0062036 | 0.0017755 | -0.0044518 | 0.00055469 | 0.001137 | 0.00053891 | 0.0020862 | 0.0061429 | -0.00012749 | 0.0016159 | 0.0039615 | -0.00023344 | 0.0040661 | 0.00028781 | 0.00049595 | 0.0013233 | -0.0073754 | 0.018475
da2 | -1.2683 | 10.452 | 0.0032555 | -0.071088 | 0.03641 | 0.14818 | 0.16888 | 0.69602 | 0.78035 | -0.11829 | 0.37417 | 0.14169 | -1.3888 | -0.43281 | -0.14364 | -0.27307 | -0.25616 | -0.39817 | 0.41028 | 0.32278 | -9.2324
da3 | -0.41462 | -3.5169 | -0.82676 | 0.22133 | -0.35595 | -0.0092415 | -0.12565 | 0.32913 | 0.26813 | 0.078851 | -0.24529 | 0.25258 | 0.20503 | 0.2877 | -0.23754 | 0.15414 | 0.073538 | -0.050957 | 0.13716 | -0.15621 | -0.094723
da4 | -0.13551 | 2.4934 | 0.17417 | 0.020566 | 0.0033463 | 0.035508 | -0.026038 | 0.079904 | 0.26838 | 0.1207 | -0.33218 | -0.21769 | -0.085158 | -0.12426 | -0.087333 | 0.035661 | 0.0096818 | -0.0036772 | -0.007396 | 0.13582 | -1.0106
da5 | -0.12417 | 1.6315 | -0.30904 | -0.10988 | -0.12341 | 0.049549 | 0.037077 | 0.1606 | 0.26771 | 0.20708 | -0.03021 | 0.028107 | -0.071728 | -0.043053 | -0.069641 | 0.18222 | -0.036945 | -0.10472 | 0.033018 | -0.06673 | -0.082814
da6 | -0.019996 | 0.39563 | 0.07505 | 0.0061951 | -0.005997 | 0.043473 | 0.012886 | 0.045978 | 0.081746 | -0.057096 | -0.074546 | -0.01983 | -0.022279 | -0.046137 | -0.0082342 | 0.013592 | -0.0094697 | 0.005554 | -0.055957 | 0.015071 | -0.29254
da7 | 0.00083476 | 1.3603 | -0.041846 | -0.037894 | -0.0011338 | -0.014098 | 0.046534 | 0.063254 | 0.048789 | 0.036193 | -0.010016 | -0.060014 | -0.053861 | 0.0037761 | -0.019655 | 0.07164 | 0.010948 | -0.021523 | 0.034715 | -0.055808 | 0.16916
da8 | 0.00478 | 0.0076448 | 0.0077769 | -0.01636 | -0.021788 | 0.011972 | 0.019192 | 0.0080731 | 0.048806 | -0.0092568 | -0.02735 | -0.0097228 | 0.0096093 | -0.014285 | 0.01918 | 0.0067339 | -0.010273 | 0.013054 | -0.022961 | -0.0124 | -0.010571
da9 | -0.025924 | -0.32776 | -0.055182 | -0.024673 | -0.024443 | -0.013752 | 0.016997 | 0.021705 | 0.033444 | 0.0034446 | -0.0069932 | 0.037506 | -0.010733 | -0.0024321 | -0.001898 | 0.029561 | 0.013416 | -0.0032589 | -0.0018977 | -0.01081 | -0.069645
da10 | 0.0055842 | 0.051749 | 0.071008 | 0.040271 | -0.0060313 | 0.029671 | 0.0079387 | -0.013425 | 0.012998 | -0.017457 | -0.068247 | -0.015579 | -0.024944 | -0.024062 | 0.0008513 | -0.0090593 | -0.0062928 | 0.0074891 | -0.023699 | 0.038569 | 0.018619
da11 | -0.048521 | 0.13482 | -0.01504 | -0.016782 | -0.0087252 | -0.00033247 | 0.0075951 | 0.0084161 | 0.0060357 | 0.0069851 | 0.00058542 | 0.0030642 | -0.0063893 | 0.0023439 | 0.0067391 | 0.014795 | 0.00030343 | -0.0031388 | 0.0022092 | -0.0086645 | -0.024144
da12 | -0.0028263 | -0.010006 | 0.011537 | 0.0040804 | -0.0015246 | 0.0062151 | 0.0012842 | 0.0017974 | 0.0037311 | -0.004283 | -0.0089779 | -0.00040592 | -0.00032249 | -0.0053053 | -0.0010287 | -0.001726 | -0.0014925 | -0.00033569 | -0.0069124 | 0.0036694 | -0.01095
da13 | -0.0069414 | -0.023261 | -0.0070297 | -0.0014539 | -0.0029128 | 0.0011151 | 0.00079323 | 0.0031626 | 0.0013895 | 0.00075044 | 0.0010528 | 0.0036291 | -0.001171 | -0.00099604 | 0.00013965 | 0.0034931 | 0.0013559 | -0.0015633 | 5.1917e-05 | -0.0018066 | -0.0025784
da14 | -0.022212 | 0.0097799 | 0.0037739 | 0.0032259 | -0.0020155 | 0.0022 | 0.0013026 | 0.00078267 | 0.0014505 | -0.0010659 | -0.0049058 | -0.00034734 | -0.003021 | -0.0032409 | 6.6725e-05 | -0.0017364 | -0.00065468 | 0.00049775 | -0.00097682 | 0.0046645 | -0.00045497
da15 | -0.016506 | 0.0086944 | 0.0036192 | -0.0017828 | 0.0032068 | -0.0017614 | 0.0059493 | -0.0053304 | 0.010866 | -0.0023572 | -0.010239 | -0.0061211 | -0.0034223 | 0.011532 | -0.0021825 | 0.00049394 | -0.0034599 | 0.0041813 | -0.0039503 | 0.00075822 | 0.011409
Check Computation for Aperture2
Magnetic Length
mbh dB/B | -3674.2391
| mean 3:18 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
dB/B | 594.5809 | -3527.8187 | 6.0596 | 0.12367 | 0.12367 | -1.855 | 0.12367 | -1.855 | -1.855 | -1.855 | -1.855 | -1.855 | -1.855 | -1.855 | 0.12367 | 4.081 | 4.081 | 4.081 | 2.1023 | 6.0596 | -3820.6596
| mean 2:19 | position 1 | position 2 | position 3 | position 4 | position 5 | position 6 | position 7 | position 8 | position 9 | position 10 | position 11 | position 12 | position 13 | position 14 | position 15 | position 16 | position 17 | position 18 | position 19 | position 20 | integral
D Angle | -0.088874 | 4.4216 | -0.59701 | -0.37984 | 0.071665 | 0.9045 | 0.19069 | -0.41117 | -0.28963 | 0.14979 | 0.43175 | 0.53921 | 0.20071 | 0.16312 | 0.16382 | 0.15306 | -0.20713 | -0.42008 | -0.36191 | -0.30154 | -1.67
db2 | 0.32964 | -3.9627 | 0.71113 | 0.96795 | -0.087176 | 0.76121 | 0.49277 | 1.0305 | 0.83247 | -0.53517 | -0.82409 | -0.50802 | -0.027441 | -0.057886 | -0.79151 | 0.40651 | -0.083032 | -0.77861 | -0.90925 | -0.60035 | 0.9101
db3 | -4.379 | 17.4317 | -0.2173 | -0.24492 | 0.27199 | -0.30992 | -0.45858 | -0.38352 | -0.26073 | 0.049954 | 0.11945 | 0.063811 | 0.21357 | 0.56443 | 0.28048 | 0.062729 | 0.43359 | 0.32556 | -0.27784 | -0.23275 | -11.3908
db4 | -0.014767 | -0.83983 | 0.031501 | 0.24804 | 0.1567 | -0.02446 | -0.051603 | -0.0091379 | -0.16159 | -0.10366 | -0.069871 | 0.072993 | -0.070018 | 0.0051781 | -0.098786 | -0.18038 | 0.11096 | 0.00055646 | 0.027809 | 0.11577 | 0.27454
db5 | 0.4072 | -7.1816 | -0.16394 | -0.34481 | -0.15192 | -0.052224 | 0.17009 | 0.17267 | 0.074509 | 0.061223 | 0.11069 | 0.11023 | 0.035548 | 0.057203 | 0.078796 | 0.11896 | 0.061246 | -0.067497 | -0.065481 | -0.20529 | -4.9477
db6 | -0.023376 | -0.29512 | 0.031552 | 0.055198 | 0.11126 | 0.03776 | -0.073061 | 0.021703 | 0.02879 | 0.059454 | -0.048063 | 0.095449 | -0.075712 | -0.0095771 | -0.17128 | 0.005227 | -0.042207 | 0.0073406 | -0.07188 | 0.038046 | -0.0047589
db7 | 0.9568 | 1.6539 | -0.09128 | -0.048375 | -0.02085 | -0.0044067 | -0.054525 | 0.063834 | 0.049782 | 0.048028 | -0.0010582 | -0.017928 | 0.051067 | 0.026024 | 0.059653 | -0.047467 | 0.02367 | -0.029891 | 0.0043777 | -0.010656 | 0.13699
db8 | -0.0032841 | -0.00076641 | -0.010933 | -0.030518 | 0.063373 | 0.027044 | -0.047368 | 0.032971 | 0.016306 | 0.029019 | -0.025051 | 0.055724 | -0.049475 | -0.00046001 | -0.053506 | 0.0027677 | -0.03522 | 0.051357 | -0.034799 | 0.0087687 | -0.024742
db9 | 0.58007 | 0.21848 | -0.0057549 | -0.012275 | 0.0055497 | 0.0046787 | -0.0050287 | 0.0057303 | -0.00053567 | 0.00091943 | -0.0048615 | 0.008682 | 0.0066263 | 0.003757 | -0.0013719 | -0.0084485 | 0.0032791 | 0.0061623 | -0.0024283 | -0.0046809 | 0.34177
db10 | 0.0058399 | -0.023583 | 0.032747 | -0.047026 | 0.10704 | 0.023224 | -0.077321 | 0.059847 | 0.047693 | 0.045482 | -0.061582 | 0.092314 | -0.072019 | 0.019866 | -0.086794 | -0.0005225 | -0.074056 | 0.076707 | -0.072727 | -0.012878 | -0.034806
db11 | 0.74088 | 0.58855 | -0.0038017 | -0.0053433 | -0.00034376 | 0.001444 | 0.00096114 | 0.0022436 | -0.0035782 | -0.0068466 | -0.0020517 | 0.00071964 | -0.0060155 | -0.0028156 | -0.0045331 | 0.013947 | 0.0084868 | 0.011587 | 0.00070174 | -0.0047613 | 0.57509
db12 | 0.0018193 | 0.011344 | 0.0021571 | -0.0088141 | 0.014462 | 0.0046967 | -0.012199 | 0.0098319 | 0.0081057 | 0.0041082 | -0.0080655 | 0.011255 | -0.0081611 | 0.0030642 | -0.0089261 | 0.0038125 | -0.01095 | 0.012143 | -0.011795 | -0.0047259 | -0.0056653
db13 | 0.0808 | 0.069848 | -0.00084961 | -0.00096156 | 0.0029501 | -0.0024783 | 7.0126e-05 | -0.0025123 | -0.0019437 | -0.0007924 | -0.0009834 | 0.00094456 | 0.00094092 | -0.0014936 | 0.0010451 | 0.0025281 | 0.0027243 | 0.0036966 | -0.0024049 | -0.00047998 | 0.045883
db14 | -0.0070393 | 0.0059332 | 0.0010051 | -0.0040666 | 0.0081188 | 0.0034374 | -0.0055894 | 0.0041804 | 0.0046188 | 0.0045332 | -0.0048832 | 0.0088888 | -0.0060555 | 0.0020721 | -0.0057562 | 0.0018462 | -0.0057123 | 0.0059442 | -0.0093498 | -0.0032321 | -0.0086721
db15 | 0.046441 | 0.0051604 | 0.0030485 | -0.0018459 | -0.0041776 | 0.0053175 | -0.003035 | 0.0067642 | -0.00077053 | 0.003772 | 0.0026193 | -0.0042039 | -0.0015338 | 0.0014858 | -0.0061972 | -0.0019866 | -0.005041 | -0.0028355 | 0.0053788 | 0.003241 | 0.032659
da2 | 0.80148 | 6.5221 | 0.32577 | 0.23543 | 0.48959 | 0.12035 | 1.0634 | 0.57002 | -0.25582 | -0.11974 | 0.20108 | -0.23549 | -0.16309 | 0.061787 | -0.060912 | -0.087317 | -0.29623 | -0.71844 | -0.82906 | -0.30129 | -3.6779
da3 | -0.20663 | -1.6778 | -0.17955 | 0.2787 | 0.14122 | 0.584 | 0.19895 | -0.22299 | -0.35184 | 0.13711 | -0.6819 | 0.30332 | -0.010951 | 0.019743 | 0.31698 | 0.19615 | -0.051645 | -0.092899 | -0.21827 | -0.36613 | 0.31898
da4 | -0.0053034 | 1.2381 | -0.13402 | -0.35261 | 0.015817 | 0.03694 | 0.16387 | -0.11263 | -0.067108 | 0.22554 | 0.09476 | 0.14356 | 0.23886 | 0.19601 | 0.11755 | 0.026461 | -0.14991 | -0.20303 | -0.26615 | 0.026101 | 0.2744
da5 | -0.029354 | 2.0464 | -0.23913 | -0.0088835 | -0.031832 | 0.10941 | -0.033543 | -0.016974 | 0.077165 | 0.076896 | -0.078867 | 0.0035483 | -0.06308 | -0.03864 | 0.052668 | 0.18914 | 0.00044165 | 0.020586 | 0.016209 | -0.035125 | -0.048267
da6 | 0.024053 | -0.061993 | -0.044952 | -0.076459 | 0.074828 | -0.025895 | 0.065426 | -0.067355 | 0.00057702 | 0.06183 | 0.067052 | 0.0038856 | 0.028085 | -0.031003 | 0.0262 | -0.033165 | 0.021648 | 9.9089e-05 | -0.056479 | -0.014323 | -0.15684
da7 | 0.025748 | 1.5727 | -0.093993 | -0.042881 | -0.044745 | 0.011597 | 0.034498 | 0.023977 | 0.014478 | 0.12181 | 0.025916 | 0.0041331 | -0.028388 | -0.031279 | -0.0092 | 0.095755 | -0.0019307 | -0.0035695 | 0.0098272 | -0.086007 | 0.051442
da8 | 0.02793 | -0.015878 | 0.0016703 | -0.015143 | 0.064425 | 0.00077762 | 0.035237 | -0.027588 | -0.02261 | -0.00019292 | -0.00050659 | 0.015559 | -0.014714 | -0.010368 | 0.025863 | 0.0041698 | 0.012499 | 0.015828 | -0.051448 | -0.033457 | -0.017324
da9 | -0.014498 | -0.2666 | -0.024524 | -0.0020222 | -0.016637 | -0.022145 | 0.028998 | 0.0075389 | 0.023284 | -0.0044062 | -0.001949 | 0.030819 | -0.0099778 | 0.00066114 | -0.023495 | 0.020572 | 0.0021893 | 0.014214 | -0.00091642 | -0.022202 | -0.041126
da10 | 0.0097564 | -0.035499 | -0.0085963 | -0.017814 | 0.056671 | -0.031893 | 0.015903 | -0.043374 | 0.016353 | -0.020297 | -0.016249 | 0.040442 | -0.013593 | -0.016601 | 0.024481 | 0.021413 | 0.042675 | 0.04742 | -0.063346 | -0.033596 | -0.046036
da11 | -0.043309 | 0.13708 | -0.023087 | -0.010254 | -0.003822 | 0.0033439 | 0.0094523 | 0.0024992 | 0.0042508 | 0.014974 | 0.0053585 | 0.0066436 | -0.0024805 | -0.001926 | -0.00069319 | 0.010859 | 0.0020196 | 0.0010925 | -0.0045439 | -0.013687 | -0.033023
da12 | 0.00072884 | -0.0050551 | -0.0026787 | -0.0030616 | 0.0088064 | -0.0041426 | 0.0056067 | -0.0078977 | 0.0010827 | -0.0016889 | 0.0021803 | 0.0053152 | -0.00088164 | -0.0038778 | 0.0034244 | -0.00059517 | 0.005403 | 0.0051777 | -0.0095302 | -0.002642 | -0.0039296
da13 | -0.0054911 | -0.034786 | -0.0039844 | 0.00056627 | -0.0039792 | -0.00056107 | 0.0032669 | -0.0003209 | -0.00047101 | -0.00028363 | 0.0010167 | 0.0001743 | 6.5224e-05 | 0.0014316 | 0.0012112 | 0.0022403 | 0.0019952 | -0.0018503 | 0.0019089 | -0.002426 | -0.0046385
da14 | -0.021535 | 0.0033077 | -0.00042019 | -0.0018191 | 0.0046041 | -0.0047547 | 0.0033719 | -0.0058421 | -0.00035213 | -0.0021585 | -0.00037963 | 0.0024912 | 0.0011561 | 0.00021176 | 0.0049091 | 0.0018625 | 0.0027161 | 0.0015608 | -0.0048456 | -0.0023116 | -0.0057246
da15 | -0.012786 | -0.0078717 | 0.0012392 | -0.0073764 | 0.011004 | 0.0043791 | -0.0081304 | 0.0059178 | 0.0054313 | 0.0062459 | -0.0075425 | 0.011851 | -0.0084506 | 0.0016773 | -0.008059 | 0.0021135 | -0.0063851 | 0.0092675 | -0.010234 | -0.0029492 | -0.01741
Field Colinearity | -0.021982
b1 |  | b3 |  | b5 |  | b7 |  | b9 |  | b11
inner | outer | inner | outer | inner | outer | inner | outer | inner | outer | inner | outer
59 | 40.2 | 21.8 | 16.2 | -4 | -0.77 | 1.5 | -0.22 | -0.56 | 0.031 | 0.083 | 0
Sensitivity to shims

## Alstom Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | 14.45 | 0.01 | 3.5 | 7 | Field Colinearity |  | 0 | 0.32 | 3.5 | 7
dB/B Heads CS+NCS (units) | -3750 | 130 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 595.85 | 0.36 | 4 | 8 | 0 | 1.5 | 4 | 8 | -3848.5 | 151 | 3.5 | 7 | -3657.5 | 284 | 3.5 | 7 | 8.2 | 2 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | 0 | 0.5 | 4 | 8 | 5 | 2 | 3.5 | 7 | -0.4 | 1 | 3.5 | 7
b2 (units) | 0 | 0.64 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 0 | 5 | 3.5 | 7 | 0 | 2 | 3.5 | 7
b3 (units) | 0 | 2 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 40.6 | 4.6 | 3.5 | 7 | -3.47 | 2.5 | 3.5 | 7
b4 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | 0 | 0.9 | 3.5 | 7 | 0 | 0.38 | 3.5 | 7
b5 (units) | 1.34 | 0.5 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | -2.45 | 1.3 | 3.5 | 7 | -3 | 0.7 | 3.5 | 7
b6 (units) | 0 | 0.054 | 3.5 | 7 | 0 | 0.07 | 4 | 8 | 0 | 0.34 | 3.5 | 7 | 0 | 0.1 | 3.5 | 7
b7 (units) | 0.73 | 0.09 | 3.5 | 7 | 0 | 0.054 | 4 | 8 | 2.61 | 0.3 | 3.5 | 7 | 0.62 | 0.2 | 3.5 | 7
b8 (units) | 0 | 0.02 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.05 | 3.5 | 7
b9 (units) | 0.33 | 0.03 | 3.5 | 7 | 0 | 0.02 | 4 | 8 | 0.266 | 0.05 | 3.5 | 7 | 0.23 | 0.06 | 3.5 | 7
b10 (units) | 0 | 0.0045 | 3.5 | 7 | 0 | 0.04 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.075 | 3.5 | 7
b11 (units) | 0.74 | 0.011 | 3.5 | 7 | 0 | 0.0059 | 4 | 8 | 0.632 | 0.03 | 3.5 | 7 | 0.65 | 0.018 | 3.5 | 7
b12 (units) | 0 | 0.0018 | 3.5 | 7 | 0 | 0.0061 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.01 | 3.5 | 7
b13 (units) | 0.085 | 0.0083 | 3.5 | 7 | 0 | 0.0025 | 4 | 8 | 0.0924 | 0.01 | 3.5 | 7 | 0.059 | 0.009 | 3.5 | 7
b14 (units) | 0 | 0.006 | 3.5 | 7 | 0 | 0.0034 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | 0 | 0.0035 | 3.5 | 7
b15 (units) | 0.034 | 0.007 | 3.5 | 7 | 0 | 0.0027 | 4 | 8 | -0.0018 | 0.006 | 3.5 | 7 | 0.0129 | 0.0068 | 3.5 | 7
a2 (units) | 0 | 1 | 3.5 | 7 | 0 | 1.1 | 4 | 8 | 0 | 6 | 3.5 | 7 | 0 | 3 | 3.5 | 7
a3 (units) | -0.13 | 0.35 | 3.5 | 7 | 0 | 0.35 | 4 | 8 | -1.7 | 2.1 | 3.5 | 7 | 1.35 | 0.7 | 3.5 | 7
a4 (units) | 0 | 0.27 | 3.5 | 7 | 0 | 0.27 | 4 | 8 | -0.853 | 1.6 | 3.5 | 7 | 0 | 0.7 | 3.5 | 7
a5 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.12 | 4 | 8 | 2.4 | 0.83 | 3.5 | 7 | -0.18 | 0.25 | 3.5 | 7
a6 (units) | 0 | 0.12 | 3.5 | 7 | 0 | 0.068 | 4 | 8 | 0 | 0.4 | 3.5 | 7 | 0 | 0.17 | 3.5 | 7
a7 (units) | 0.005 | 0.03 | 3.5 | 7 | 0 | 0.056 | 4 | 8 | 1.9 | 0.25 | 3.5 | 7 | -0.0102 | 0.077 | 3.5 | 7
a8 (units) | -0.02 | 0.035 | 3.5 | 7 | 0 | 0.025 | 4 | 8 | 0 | 0.12 | 3.5 | 7 | 0 | 0.06 | 3.5 | 7
a9 (units) | 0.015 | 0.03 | 3.5 | 7 | 0 | 0.021 | 4 | 8 | -0.114 | 0.05 | 3.5 | 7 | 0 | 0.024 | 3.5 | 7
a10 (units) | 0.002 | 0.003 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.085 | 3.5 | 7 | 0 | 0.045 | 3.5 | 7
a11 (units) | 0.0127 | 0.03 | 3.5 | 7 | 0 | 0.009 | 4 | 8 | 0.2 | 0.025 | 3.5 | 7 | 0 | 0.026 | 3.5 | 7
a12 (units) | 0 | 0.005 | 3.5 | 7 | 0 | 0.0056 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
a13 (units) | 0 | 0.004 | 3.5 | 7 | 0 | 0.0022 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.0037 | 3.5 | 7
a14 (units) | -0.012 | 0.006 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | -0.003 | 0.004 | 3.5 | 7
a15 (units) | 0 | 0.005 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Magnetic Length (mm) | 14.415 | 14.485 | 14.38 | 14.52 | Field Colinearity |  | -1.12 | 1.12 | -2.24 | 2.24
dB/B Heads CS+NCS (units) | -4205 | -3295 | -4660 | -2840
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 594.41 | 597.29 | 592.97 | 598.73 | -6 | 6 | -12 | 12 | -4377 | -3320 | -4905.5 | -2791.5 | -4651.5 | -2663.5 | -5645.5 | -1669.5 | 1.2 | 15.2 | -5.8 | 22.2
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -2 | 2 | -4 | 4 | -2 | 12 | -9 | 19 | -3.9 | 3.1 | -7.4 | 6.6
b2 (units) | -2.24 | 2.24 | -4.48 | 4.48 | -2.4 | 2.4 | -4.8 | 4.8 | -17.5 | 17.5 | -35 | 35 | -7 | 7 | -14 | 14
b3 (units) | -7 | 7 | -14 | 14 | -2.4 | 2.4 | -4.8 | 4.8 | 24.5 | 56.7 | 8.4 | 72.8 | -12.22 | 5.28 | -20.97 | 14.03
b4 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.64 | 0.64 | -1.28 | 1.28 | -3.15 | 3.15 | -6.3 | 6.3 | -1.33 | 1.33 | -2.66 | 2.66
b5 (units) | -0.41 | 3.09 | -2.16 | 4.84 | -0.64 | 0.64 | -1.28 | 1.28 | -7 | 2.1 | -11.55 | 6.65 | -5.45 | -0.55 | -7.9 | 1.9
b6 (units) | -0.189 | 0.189 | -0.378 | 0.378 | -0.28 | 0.28 | -0.56 | 0.56 | -1.19 | 1.19 | -2.38 | 2.38 | -0.35 | 0.35 | -0.7 | 0.7
b7 (units) | 0.415 | 1.045 | 0.1 | 1.36 | -0.216 | 0.216 | -0.432 | 0.432 | 1.56 | 3.66 | 0.51 | 4.71 | -0.08 | 1.32 | -0.78 | 2.02
b8 (units) | -0.07 | 0.07 | -0.14 | 0.14 | -0.136 | 0.136 | -0.272 | 0.272 | -0.35 | 0.35 | -0.7 | 0.7 | -0.175 | 0.175 | -0.35 | 0.35
b9 (units) | 0.225 | 0.435 | 0.12 | 0.54 | -0.08 | 0.08 | -0.16 | 0.16 | 0.091 | 0.441 | -0.084 | 0.616 | 0.02 | 0.44 | -0.19 | 0.65
b10 (units) | -0.01575 | 0.01575 | -0.0315 | 0.0315 | -0.16 | 0.16 | -0.32 | 0.32 | -0.35 | 0.35 | -0.7 | 0.7 | -0.2625 | 0.2625 | -0.525 | 0.525
b11 (units) | 0.7015 | 0.7785 | 0.663 | 0.817 | -0.0236 | 0.0236 | -0.0472 | 0.0472 | 0.527 | 0.737 | 0.422 | 0.842 | 0.587 | 0.713 | 0.524 | 0.776
b12 (units) | -0.0063 | 0.0063 | -0.0126 | 0.0126 | -0.0244 | 0.0244 | -0.0488 | 0.0488 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.035 | 0.035 | -0.07 | 0.07
b13 (units) | 0.05595 | 0.11405 | 0.0269 | 0.1431 | -0.01 | 0.01 | -0.02 | 0.02 | 0.0574 | 0.1274 | 0.0224 | 0.1624 | 0.0275 | 0.0905 | -0.004 | 0.122
b14 (units) | -0.021 | 0.021 | -0.042 | 0.042 | -0.0136 | 0.0136 | -0.0272 | 0.0272 | -0.021 | 0.021 | -0.042 | 0.042 | -0.01225 | 0.01225 | -0.0245 | 0.0245
b15 (units) | 0.0095 | 0.0585 | -0.015 | 0.083 | -0.0108 | 0.0108 | -0.0216 | 0.0216 | -0.0228 | 0.0192 | -0.0438 | 0.0402 | -0.0109 | 0.0367 | -0.0347 | 0.0605
a2 (units) | -3.5 | 3.5 | -7 | 7 | -4.4 | 4.4 | -8.8 | 8.8 | -21 | 21 | -42 | 42 | -10.5 | 10.5 | -21 | 21
a3 (units) | -1.355 | 1.095 | -2.58 | 2.32 | -1.4 | 1.4 | -2.8 | 2.8 | -9.05 | 5.65 | -16.4 | 13 | -1.1 | 3.8 | -3.55 | 6.25
a4 (units) | -0.945 | 0.945 | -1.89 | 1.89 | -1.08 | 1.08 | -2.16 | 2.16 | -6.453 | 4.747 | -12.053 | 10.347 | -2.45 | 2.45 | -4.9 | 4.9
a5 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.48 | 0.48 | -0.96 | 0.96 | -0.505 | 5.305 | -3.41 | 8.21 | -1.055 | 0.695 | -1.93 | 1.57
a6 (units) | -0.42 | 0.42 | -0.84 | 0.84 | -0.272 | 0.272 | -0.544 | 0.544 | -1.4 | 1.4 | -2.8 | 2.8 | -0.595 | 0.595 | -1.19 | 1.19
a7 (units) | -0.1 | 0.11 | -0.205 | 0.215 | -0.224 | 0.224 | -0.448 | 0.448 | 1.025 | 2.775 | 0.15 | 3.65 | -0.2797 | 0.2593 | -0.5492 | 0.5288
a8 (units) | -0.1425 | 0.1025 | -0.265 | 0.225 | -0.1 | 0.1 | -0.2 | 0.2 | -0.42 | 0.42 | -0.84 | 0.84 | -0.21 | 0.21 | -0.42 | 0.42
a9 (units) | -0.09 | 0.12 | -0.195 | 0.225 | -0.084 | 0.084 | -0.168 | 0.168 | -0.289 | 0.061 | -0.464 | 0.236 | -0.084 | 0.084 | -0.168 | 0.168
a10 (units) | -0.0085 | 0.0125 | -0.019 | 0.023 | -0.136 | 0.136 | -0.272 | 0.272 | -0.2975 | 0.2975 | -0.595 | 0.595 | -0.1575 | 0.1575 | -0.315 | 0.315
a11 (units) | -0.0923 | 0.1177 | -0.1973 | 0.2227 | -0.036 | 0.036 | -0.072 | 0.072 | 0.1125 | 0.2875 | 0.025 | 0.375 | -0.091 | 0.091 | -0.182 | 0.182
a12 (units) | -0.0175 | 0.0175 | -0.035 | 0.035 | -0.0224 | 0.0224 | -0.0448 | 0.0448 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.021 | 0.021 | -0.042 | 0.042
a13 (units) | -0.014 | 0.014 | -0.028 | 0.028 | -0.0088 | 0.0088 | -0.0176 | 0.0176 | -0.035 | 0.035 | -0.07 | 0.07 | -0.01295 | 0.01295 | -0.0259 | 0.0259
a14 (units) | -0.033 | 0.009 | -0.054 | 0.03 | -0.012 | 0.012 | -0.024 | 0.024 | -0.021 | 0.021 | -0.042 | 0.042 | -0.017 | 0.011 | -0.031 | 0.025
a15 (units) | -0.0175 | 0.0175 | -0.035 | 0.035 | -0.012 | 0.012 | -0.024 | 0.024 | -0.035 | 0.035 | -0.07 | 0.07 | -0.021 | 0.021 | -0.042 | 0.042

## Ansaldo Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | 14.45 | 0.01 | 3.5 | 7 | Field Colinearity |  | 0 | 0.32 | 3.5 | 7
dB/B Heads CS+NCS (units) | -3750 | 130 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 595.85 | 0.36 | 4 | 8 | 0 | 1.5 | 4 | 8 | -3848.5 | 151 | 3.5 | 7 | -3657.5 | 284 | 3.5 | 7 | 8.2 | 2 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | 0 | 0.5 | 4 | 8 | 5 | 2 | 3.5 | 7 | -0.4 | 1 | 3.5 | 7
b2 (units) | 0 | 0.64 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 0 | 5 | 3.5 | 7 | 0 | 2 | 3.5 | 7
b3 (units) | 0 | 2 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 24 | 4.6 | 3.5 | 7 | -6.652 | 2.5 | 3.5 | 7
b4 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | 0 | 0.9 | 3.5 | 7 | 0 | 0.38 | 3.5 | 7
b5 (units) | 1.43 | 0.5 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | -5.57 | 1.3 | 3.5 | 7 | -3 | 0.7 | 3.5 | 7
b6 (units) | 0 | 0.054 | 3.5 | 7 | 0 | 0.07 | 4 | 8 | 0 | 0.34 | 3.5 | 7 | 0 | 0.1 | 3.5 | 7
b7 (units) | 0.54 | 0.09 | 3.5 | 7 | 0 | 0.054 | 4 | 8 | 1.62 | 0.3 | 3.5 | 7 | -0.0495 | 0.2 | 3.5 | 7
b8 (units) | 0 | 0.02 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.05 | 3.5 | 7
b9 (units) | 0.28 | 0.03 | 3.5 | 7 | 0 | 0.02 | 4 | 8 | 0.147 | 0.05 | 3.5 | 7 | 0.23 | 0.06 | 3.5 | 7
b10 (units) | 0 | 0.0045 | 3.5 | 7 | 0 | 0.04 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.075 | 3.5 | 7
b11 (units) | 0.74 | 0.011 | 3.5 | 7 | 0 | 0.0059 | 4 | 8 | 0.55 | 0.03 | 3.5 | 7 | 0.65 | 0.018 | 3.5 | 7
b12 (units) | 0 | 0.0018 | 3.5 | 7 | 0 | 0.0061 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.01 | 3.5 | 7
b13 (units) | 0.085 | 0.0083 | 3.5 | 7 | 0 | 0.0025 | 4 | 8 | 0.0807 | 0.01 | 3.5 | 7 | 0.059 | 0.009 | 3.5 | 7
b14 (units) | 0 | 0.006 | 3.5 | 7 | 0 | 0.0034 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | 0 | 0.0035 | 3.5 | 7
b15 (units) | 0.043 | 0.007 | 3.5 | 7 | 0 | 0.0027 | 4 | 8 | -0.00515 | 0.006 | 3.5 | 7 | 0.0244 | 0.0068 | 3.5 | 7
a2 (units) | 0 | 1 | 3.5 | 7 | 0 | 1.1 | 4 | 8 | 0 | 6 | 3.5 | 7 | 0 | 3 | 3.5 | 7
a3 (units) | -0.46 | 0.35 | 3.5 | 7 | 0 | 0.35 | 4 | 8 | -1.7 | 2.1 | 3.5 | 7 | -0.113 | 0.7 | 3.5 | 7
a4 (units) | 0 | 0.27 | 3.5 | 7 | 0 | 0.27 | 4 | 8 | 2.02 | 1.6 | 3.5 | 7 | 0 | 0.7 | 3.5 | 7
a5 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.12 | 4 | 8 | 2.4 | 0.83 | 3.5 | 7 | -0.18 | 0.25 | 3.5 | 7
a6 (units) | 0 | 0.12 | 3.5 | 7 | 0 | 0.068 | 4 | 8 | 0 | 0.4 | 3.5 | 7 | 0 | 0.17 | 3.5 | 7
a7 (units) | 0.044 | 0.03 | 3.5 | 7 | 0 | 0.056 | 4 | 8 | 1.9 | 0.25 | 3.5 | 7 | 0.108 | 0.077 | 3.5 | 7
a8 (units) | -0.02 | 0.035 | 3.5 | 7 | 0 | 0.025 | 4 | 8 | 0 | 0.12 | 3.5 | 7 | 0 | 0.06 | 3.5 | 7
a9 (units) | 0.035 | 0.03 | 3.5 | 7 | 0 | 0.021 | 4 | 8 | -0.243 | 0.05 | 3.5 | 7 | 0 | 0.024 | 3.5 | 7
a10 (units) | 0.002 | 0.003 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.085 | 3.5 | 7 | 0 | 0.045 | 3.5 | 7
a11 (units) | 0.037 | 0.03 | 3.5 | 7 | 0 | 0.009 | 4 | 8 | 0.2 | 0.025 | 3.5 | 7 | 0 | 0.026 | 3.5 | 7
a12 (units) | 0 | 0.005 | 3.5 | 7 | 0 | 0.0056 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
a13 (units) | 0 | 0.004 | 3.5 | 7 | 0 | 0.0022 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.0037 | 3.5 | 7
a14 (units) | -0.012 | 0.006 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | -0.003 | 0.004 | 3.5 | 7
a15 (units) | 0.0058 | 0.005 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Magnetic Length (mm) | 14.415 | 14.485 | 14.38 | 14.52 | Field Colinearity |  | -1.12 | 1.12 | -2.24 | 2.24
dB/B Heads CS+NCS (units) | -4205 | -3295 | -4660 | -2840
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 594.41 | 597.29 | 592.97 | 598.73 | -6 | 6 | -12 | 12 | -4377 | -3320 | -4905.5 | -2791.5 | -4651.5 | -2663.5 | -5645.5 | -1669.5 | 1.2 | 15.2 | -5.8 | 22.2
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -2 | 2 | -4 | 4 | -2 | 12 | -9 | 19 | -3.9 | 3.1 | -7.4 | 6.6
b2 (units) | -2.24 | 2.24 | -4.48 | 4.48 | -2.4 | 2.4 | -4.8 | 4.8 | -17.5 | 17.5 | -35 | 35 | -7 | 7 | -14 | 14
b3 (units) | -7 | 7 | -14 | 14 | -2.4 | 2.4 | -4.8 | 4.8 | 7.9 | 40.1 | -8.2 | 56.2 | -15.402 | 2.098 | -24.152 | 10.848
b4 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.64 | 0.64 | -1.28 | 1.28 | -3.15 | 3.15 | -6.3 | 6.3 | -1.33 | 1.33 | -2.66 | 2.66
b5 (units) | -0.32 | 3.18 | -2.07 | 4.93 | -0.64 | 0.64 | -1.28 | 1.28 | -10.12 | -1.02 | -14.67 | 3.53 | -5.45 | -0.55 | -7.9 | 1.9
b6 (units) | -0.189 | 0.189 | -0.378 | 0.378 | -0.28 | 0.28 | -0.56 | 0.56 | -1.19 | 1.19 | -2.38 | 2.38 | -0.35 | 0.35 | -0.7 | 0.7
b7 (units) | 0.225 | 0.855 | -0.09 | 1.17 | -0.216 | 0.216 | -0.432 | 0.432 | 0.57 | 2.67 | -0.48 | 3.72 | -0.7495 | 0.6505 | -1.4495 | 1.3505
b8 (units) | -0.07 | 0.07 | -0.14 | 0.14 | -0.136 | 0.136 | -0.272 | 0.272 | -0.35 | 0.35 | -0.7 | 0.7 | -0.175 | 0.175 | -0.35 | 0.35
b9 (units) | 0.175 | 0.385 | 0.07 | 0.49 | -0.08 | 0.08 | -0.16 | 0.16 | -0.028 | 0.322 | -0.203 | 0.497 | 0.02 | 0.44 | -0.19 | 0.65
b10 (units) | -0.01575 | 0.01575 | -0.0315 | 0.0315 | -0.16 | 0.16 | -0.32 | 0.32 | -0.35 | 0.35 | -0.7 | 0.7 | -0.2625 | 0.2625 | -0.525 | 0.525
b11 (units) | 0.7015 | 0.7785 | 0.663 | 0.817 | -0.0236 | 0.0236 | -0.0472 | 0.0472 | 0.445 | 0.655 | 0.34 | 0.76 | 0.587 | 0.713 | 0.524 | 0.776
b12 (units) | -0.0063 | 0.0063 | -0.0126 | 0.0126 | -0.0244 | 0.0244 | -0.0488 | 0.0488 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.035 | 0.035 | -0.07 | 0.07
b13 (units) | 0.05595 | 0.11405 | 0.0269 | 0.1431 | -0.01 | 0.01 | -0.02 | 0.02 | 0.0457 | 0.1157 | 0.0107 | 0.1507 | 0.0275 | 0.0905 | -0.004 | 0.122
b14 (units) | -0.021 | 0.021 | -0.042 | 0.042 | -0.0136 | 0.0136 | -0.0272 | 0.0272 | -0.021 | 0.021 | -0.042 | 0.042 | -0.01225 | 0.01225 | -0.0245 | 0.0245
b15 (units) | 0.0185 | 0.0675 | -0.006 | 0.092 | -0.0108 | 0.0108 | -0.0216 | 0.0216 | -0.02615 | 0.01585 | -0.04715 | 0.03685 | 0.0006 | 0.0482 | -0.0232 | 0.072
a2 (units) | -3.5 | 3.5 | -7 | 7 | -4.4 | 4.4 | -8.8 | 8.8 | -21 | 21 | -42 | 42 | -10.5 | 10.5 | -21 | 21
a3 (units) | -1.685 | 0.765 | -2.91 | 1.99 | -1.4 | 1.4 | -2.8 | 2.8 | -9.05 | 5.65 | -16.4 | 13 | -2.563 | 2.337 | -5.013 | 4.787
a4 (units) | -0.945 | 0.945 | -1.89 | 1.89 | -1.08 | 1.08 | -2.16 | 2.16 | -3.58 | 7.62 | -9.18 | 13.22 | -2.45 | 2.45 | -4.9 | 4.9
a5 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.48 | 0.48 | -0.96 | 0.96 | -0.505 | 5.305 | -3.41 | 8.21 | -1.055 | 0.695 | -1.93 | 1.57
a6 (units) | -0.42 | 0.42 | -0.84 | 0.84 | -0.272 | 0.272 | -0.544 | 0.544 | -1.4 | 1.4 | -2.8 | 2.8 | -0.595 | 0.595 | -1.19 | 1.19
a7 (units) | -0.061 | 0.149 | -0.166 | 0.254 | -0.224 | 0.224 | -0.448 | 0.448 | 1.025 | 2.775 | 0.15 | 3.65 | -0.1615 | 0.3775 | -0.431 | 0.647
a8 (units) | -0.1425 | 0.1025 | -0.265 | 0.225 | -0.1 | 0.1 | -0.2 | 0.2 | -0.42 | 0.42 | -0.84 | 0.84 | -0.21 | 0.21 | -0.42 | 0.42
a9 (units) | -0.07 | 0.14 | -0.175 | 0.245 | -0.084 | 0.084 | -0.168 | 0.168 | -0.418 | -0.068 | -0.593 | 0.107 | -0.084 | 0.084 | -0.168 | 0.168
a10 (units) | -0.0085 | 0.0125 | -0.019 | 0.023 | -0.136 | 0.136 | -0.272 | 0.272 | -0.2975 | 0.2975 | -0.595 | 0.595 | -0.1575 | 0.1575 | -0.315 | 0.315
a11 (units) | -0.068 | 0.142 | -0.173 | 0.247 | -0.036 | 0.036 | -0.072 | 0.072 | 0.1125 | 0.2875 | 0.025 | 0.375 | -0.091 | 0.091 | -0.182 | 0.182
a12 (units) | -0.0175 | 0.0175 | -0.035 | 0.035 | -0.0224 | 0.0224 | -0.0448 | 0.0448 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.021 | 0.021 | -0.042 | 0.042
a13 (units) | -0.014 | 0.014 | -0.028 | 0.028 | -0.0088 | 0.0088 | -0.0176 | 0.0176 | -0.035 | 0.035 | -0.07 | 0.07 | -0.01295 | 0.01295 | -0.0259 | 0.0259
a14 (units) | -0.033 | 0.009 | -0.054 | 0.03 | -0.012 | 0.012 | -0.024 | 0.024 | -0.021 | 0.021 | -0.042 | 0.042 | -0.017 | 0.011 | -0.031 | 0.025
a15 (units) | -0.0117 | 0.0233 | -0.0292 | 0.0408 | -0.012 | 0.012 | -0.024 | 0.024 | -0.035 | 0.035 | -0.07 | 0.07 | -0.021 | 0.021 | -0.042 | 0.042

## Noell Bound
Bound Parameters
| Mean | Sigma | Ybound | Rbound |  |  | Mean | Sigma | Ybound | Rbound
Magnetic Length (mm) | 14.45 | 0.01 | 3.5 | 7 | Field Colinearity |  | 0 | 0.32 | 3.5 | 7
dB/B Heads CS+NCS (units) | -3750 | 130 | 3.5 | 7
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Main Field Component | 595.85 | 0.36 | 4 | 8 | 0 | 1.5 | 4 | 8 | -3848.5 | 151 | 3.5 | 7 | -3657.5 | 284 | 3.5 | 7 | 8.2 | 2 | 3.5 | 7
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound | Mean | Sigma | Ybound | Rbound
Delta Angle (mrad) | 0 | 0 | 3.5 | 7 | 0 | 0.5 | 4 | 8 | 5 | 2 | 3.5 | 7 | -0.4 | 1 | 3.5 | 7
b2 (units) | 0 | 0.64 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 0 | 5 | 3.5 | 7 | 0 | 2 | 3.5 | 7
b3 (units) | 0 | 2 | 3.5 | 7 | 0 | 0.6 | 4 | 8 | 34.5 | 4.6 | 3.5 | 7 | 2.43 | 2.5 | 3.5 | 7
b4 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | 0 | 0.9 | 3.5 | 7 | 0 | 0.38 | 3.5 | 7
b5 (units) | 0.74 | 0.5 | 3.5 | 7 | 0 | 0.16 | 4 | 8 | -4.44 | 1.3 | 3.5 | 7 | -3 | 0.7 | 3.5 | 7
b6 (units) | 0 | 0.054 | 3.5 | 7 | 0 | 0.07 | 4 | 8 | 0 | 0.34 | 3.5 | 7 | 0 | 0.1 | 3.5 | 7
b7 (units) | 0.7 | 0.09 | 3.5 | 7 | 0 | 0.054 | 4 | 8 | 2.54 | 0.3 | 3.5 | 7 | 0.389 | 0.2 | 3.5 | 7
b8 (units) | 0 | 0.02 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.05 | 3.5 | 7
b9 (units) | 0.26 | 0.03 | 3.5 | 7 | 0 | 0.02 | 4 | 8 | 0.282 | 0.05 | 3.5 | 7 | 0.23 | 0.06 | 3.5 | 7
b10 (units) | 0 | 0.0045 | 3.5 | 7 | 0 | 0.04 | 4 | 8 | 0 | 0.1 | 3.5 | 7 | 0 | 0.075 | 3.5 | 7
b11 (units) | 0.74 | 0.011 | 3.5 | 7 | 0 | 0.0059 | 4 | 8 | 0.603 | 0.03 | 3.5 | 7 | 0.65 | 0.018 | 3.5 | 7
b12 (units) | 0 | 0.0018 | 3.5 | 7 | 0 | 0.0061 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.01 | 3.5 | 7
b13 (units) | 0.085 | 0.0083 | 3.5 | 7 | 0 | 0.0025 | 4 | 8 | 0.107 | 0.01 | 3.5 | 7 | 0.059 | 0.009 | 3.5 | 7
b14 (units) | 0 | 0.006 | 3.5 | 7 | 0 | 0.0034 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | 0 | 0.0035 | 3.5 | 7
b15 (units) | 0.032 | 0.007 | 3.5 | 7 | 0 | 0.0027 | 4 | 8 | -0.008 | 0.006 | 3.5 | 7 | 0.0135 | 0.0068 | 3.5 | 7
a2 (units) | 0 | 1 | 3.5 | 7 | 0 | 1.1 | 4 | 8 | 0 | 6 | 3.5 | 7 | 0 | 3 | 3.5 | 7
a3 (units) | 0.12 | 0.35 | 3.5 | 7 | 0 | 0.35 | 4 | 8 | -1.7 | 2.1 | 3.5 | 7 | 0.988 | 0.7 | 3.5 | 7
a4 (units) | 0 | 0.27 | 3.5 | 7 | 0 | 0.27 | 4 | 8 | 0.76 | 1.6 | 3.5 | 7 | 0 | 0.7 | 3.5 | 7
a5 (units) | 0 | 0.14 | 3.5 | 7 | 0 | 0.12 | 4 | 8 | 2.4 | 0.83 | 3.5 | 7 | -0.18 | 0.25 | 3.5 | 7
a6 (units) | 0 | 0.12 | 3.5 | 7 | 0 | 0.068 | 4 | 8 | 0 | 0.4 | 3.5 | 7 | 0 | 0.17 | 3.5 | 7
a7 (units) | 0.002 | 0.03 | 3.5 | 7 | 0 | 0.056 | 4 | 8 | 1.9 | 0.25 | 3.5 | 7 | -0.00256 | 0.077 | 3.5 | 7
a8 (units) | -0.02 | 0.035 | 3.5 | 7 | 0 | 0.025 | 4 | 8 | 0 | 0.12 | 3.5 | 7 | 0 | 0.06 | 3.5 | 7
a9 (units) | 0.0072 | 0.03 | 3.5 | 7 | 0 | 0.021 | 4 | 8 | -0.231 | 0.05 | 3.5 | 7 | 0 | 0.024 | 3.5 | 7
a10 (units) | 0.002 | 0.003 | 3.5 | 7 | 0 | 0.034 | 4 | 8 | 0 | 0.085 | 3.5 | 7 | 0 | 0.045 | 3.5 | 7
a11 (units) | 0.002 | 0.03 | 3.5 | 7 | 0 | 0.009 | 4 | 8 | 0.2 | 0.025 | 3.5 | 7 | 0 | 0.026 | 3.5 | 7
a12 (units) | 0 | 0.005 | 3.5 | 7 | 0 | 0.0056 | 4 | 8 | 0 | 0.015 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
a13 (units) | 0 | 0.004 | 3.5 | 7 | 0 | 0.0022 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.0037 | 3.5 | 7
a14 (units) | -0.012 | 0.006 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.006 | 3.5 | 7 | -0.003 | 0.004 | 3.5 | 7
a15 (units) | 0 | 0.005 | 3.5 | 7 | 0 | 0.003 | 4 | 8 | 0 | 0.01 | 3.5 | 7 | 0 | 0.006 | 3.5 | 7
Bound Limits
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit |  |  | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Magnetic Length (mm) | 14.415 | 14.485 | 14.38 | 14.52 | Field Colinearity |  | -1.12 | 1.12 | -2.24 | 2.24
dB/B Heads CS+NCS (units) | -4205 | -3295 | -4660 | -2840
| Mean Value (3:18) (mT/kA) |  |  |  | dB/B Inner Positions 3 to 18 (units) |  |  |  | dB/B Head CS (units) |  |  |  | dB/B Head NCS (units) |  |  |  | dB/B Horn positions 2 and 19 (units)
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Main Field Component | 594.41 | 597.29 | 592.97 | 598.73 | -6 | 6 | -12 | 12 | -4377 | -3320 | -4905.5 | -2791.5 | -4651.5 | -2663.5 | -5645.5 | -1669.5 | 1.2 | 15.2 | -5.8 | 22.2
| Mean Value (2:19) |  |  |  | Inner Positions 2 to 19 minus mean value |  |  |  | Head CS |  |  |  | Head NCS
| Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit | Inf. Y limit | Sup. Y limit | Inf. R limit | Sup. R limit
Delta Angle (mrad) | 0 | 0 | 0 | 0 | -2 | 2 | -4 | 4 | -2 | 12 | -9 | 19 | -3.9 | 3.1 | -7.4 | 6.6
b2 (units) | -2.24 | 2.24 | -4.48 | 4.48 | -2.4 | 2.4 | -4.8 | 4.8 | -17.5 | 17.5 | -35 | 35 | -7 | 7 | -14 | 14
b3 (units) | -7 | 7 | -14 | 14 | -2.4 | 2.4 | -4.8 | 4.8 | 18.4 | 50.6 | 2.3 | 66.7 | -6.32 | 11.18 | -15.07 | 19.93
b4 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.64 | 0.64 | -1.28 | 1.28 | -3.15 | 3.15 | -6.3 | 6.3 | -1.33 | 1.33 | -2.66 | 2.66
b5 (units) | -1.01 | 2.49 | -2.76 | 4.24 | -0.64 | 0.64 | -1.28 | 1.28 | -8.99 | 0.11 | -13.54 | 4.66 | -5.45 | -0.55 | -7.9 | 1.9
b6 (units) | -0.189 | 0.189 | -0.378 | 0.378 | -0.28 | 0.28 | -0.56 | 0.56 | -1.19 | 1.19 | -2.38 | 2.38 | -0.35 | 0.35 | -0.7 | 0.7
b7 (units) | 0.385 | 1.015 | 0.07 | 1.33 | -0.216 | 0.216 | -0.432 | 0.432 | 1.49 | 3.59 | 0.44 | 4.64 | -0.311 | 1.089 | -1.011 | 1.789
b8 (units) | -0.07 | 0.07 | -0.14 | 0.14 | -0.136 | 0.136 | -0.272 | 0.272 | -0.35 | 0.35 | -0.7 | 0.7 | -0.175 | 0.175 | -0.35 | 0.35
b9 (units) | 0.155 | 0.365 | 0.05 | 0.47 | -0.08 | 0.08 | -0.16 | 0.16 | 0.107 | 0.457 | -0.068 | 0.632 | 0.02 | 0.44 | -0.19 | 0.65
b10 (units) | -0.01575 | 0.01575 | -0.0315 | 0.0315 | -0.16 | 0.16 | -0.32 | 0.32 | -0.35 | 0.35 | -0.7 | 0.7 | -0.2625 | 0.2625 | -0.525 | 0.525
b11 (units) | 0.7015 | 0.7785 | 0.663 | 0.817 | -0.0236 | 0.0236 | -0.0472 | 0.0472 | 0.498 | 0.708 | 0.393 | 0.813 | 0.587 | 0.713 | 0.524 | 0.776
b12 (units) | -0.0063 | 0.0063 | -0.0126 | 0.0126 | -0.0244 | 0.0244 | -0.0488 | 0.0488 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.035 | 0.035 | -0.07 | 0.07
b13 (units) | 0.05595 | 0.11405 | 0.0269 | 0.1431 | -0.01 | 0.01 | -0.02 | 0.02 | 0.072 | 0.142 | 0.037 | 0.177 | 0.0275 | 0.0905 | -0.004 | 0.122
b14 (units) | -0.021 | 0.021 | -0.042 | 0.042 | -0.0136 | 0.0136 | -0.0272 | 0.0272 | -0.021 | 0.021 | -0.042 | 0.042 | -0.01225 | 0.01225 | -0.0245 | 0.0245
b15 (units) | 0.0075 | 0.0565 | -0.017 | 0.081 | -0.0108 | 0.0108 | -0.0216 | 0.0216 | -0.029 | 0.013 | -0.05 | 0.034 | -0.0103 | 0.0373 | -0.0341 | 0.0611
a2 (units) | -3.5 | 3.5 | -7 | 7 | -4.4 | 4.4 | -8.8 | 8.8 | -21 | 21 | -42 | 42 | -10.5 | 10.5 | -21 | 21
a3 (units) | -1.105 | 1.345 | -2.33 | 2.57 | -1.4 | 1.4 | -2.8 | 2.8 | -9.05 | 5.65 | -16.4 | 13 | -1.462 | 3.438 | -3.912 | 5.888
a4 (units) | -0.945 | 0.945 | -1.89 | 1.89 | -1.08 | 1.08 | -2.16 | 2.16 | -4.84 | 6.36 | -10.44 | 11.96 | -2.45 | 2.45 | -4.9 | 4.9
a5 (units) | -0.49 | 0.49 | -0.98 | 0.98 | -0.48 | 0.48 | -0.96 | 0.96 | -0.505 | 5.305 | -3.41 | 8.21 | -1.055 | 0.695 | -1.93 | 1.57
a6 (units) | -0.42 | 0.42 | -0.84 | 0.84 | -0.272 | 0.272 | -0.544 | 0.544 | -1.4 | 1.4 | -2.8 | 2.8 | -0.595 | 0.595 | -1.19 | 1.19
a7 (units) | -0.103 | 0.107 | -0.208 | 0.212 | -0.224 | 0.224 | -0.448 | 0.448 | 1.025 | 2.775 | 0.15 | 3.65 | -0.27206 | 0.26694 | -0.54156 | 0.53644
a8 (units) | -0.1425 | 0.1025 | -0.265 | 0.225 | -0.1 | 0.1 | -0.2 | 0.2 | -0.42 | 0.42 | -0.84 | 0.84 | -0.21 | 0.21 | -0.42 | 0.42
a9 (units) | -0.0978 | 0.1122 | -0.2028 | 0.2172 | -0.084 | 0.084 | -0.168 | 0.168 | -0.406 | -0.056 | -0.581 | 0.119 | -0.084 | 0.084 | -0.168 | 0.168
a10 (units) | -0.0085 | 0.0125 | -0.019 | 0.023 | -0.136 | 0.136 | -0.272 | 0.272 | -0.2975 | 0.2975 | -0.595 | 0.595 | -0.1575 | 0.1575 | -0.315 | 0.315
a11 (units) | -0.103 | 0.107 | -0.208 | 0.212 | -0.036 | 0.036 | -0.072 | 0.072 | 0.1125 | 0.2875 | 0.025 | 0.375 | -0.091 | 0.091 | -0.182 | 0.182
a12 (units) | -0.0175 | 0.0175 | -0.035 | 0.035 | -0.0224 | 0.0224 | -0.0448 | 0.0448 | -0.0525 | 0.0525 | -0.105 | 0.105 | -0.021 | 0.021 | -0.042 | 0.042
a13 (units) | -0.014 | 0.014 | -0.028 | 0.028 | -0.0088 | 0.0088 | -0.0176 | 0.0176 | -0.035 | 0.035 | -0.07 | 0.07 | -0.01295 | 0.01295 | -0.0259 | 0.0259
a14 (units) | -0.033 | 0.009 | -0.054 | 0.03 | -0.012 | 0.012 | -0.024 | 0.024 | -0.021 | 0.021 | -0.042 | 0.042 | -0.017 | 0.011 | -0.031 | 0.025
a15 (units) | -0.0175 | 0.0175 | -0.035 | 0.035 | -0.012 | 0.012 | -0.024 | 0.024 | -0.035 | 0.035 | -0.07 | 0.07 | -0.021 | 0.021 | -0.042 | 0.042

## MTF
Straight part | [21 empty] | Integral
Aperture 1 | [10 empty] | Aperture 2 | [10 empty] | Aperture 1 | [9 empty] | Aperture 2
b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Coil pos. | TF | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Coil pos. | TF | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Mag len | b2 | b3 | b4 | b5 | b7 | a1 | a2 | a3 | a4 | Mag len
-0.18374 | -4.8065 | -0.079944 | 0.21023 | 0.91984 | -1.3488 | -1.2683 | -0.41462 | -0.13551 | 16.3819 | 0.5946 | 0.32964 | -4.379 | -0.014767 | 0.4072 | 0.9568 | -0.88873 | 0.80148 | -0.20663 | -0.0053034 | 15.8784 | 0.59462 | -0.31165 | -4.3362 | -0.18073 | -0.20974 | 0.91387 | -1.4853e-14 | -1.1339 | -0.50765 | -0.076091 | 14442.684 | 0.20407 | -3.8712 | -0.033194 | -0.019473 | 0.95392 | 0 | 0.84994 | -0.23914 | 0.045413 | 14446.837
Shims
Aperture 1 | [7 empty] | Aperture 2
Upper |  |  |  | Lower |  |  |  | Upper |  |  |  | Lower
Inner layer |  | Outer layer |  | Inner layer |  | Outer layer |  | Inner layer |  | Outer layer |  | Inner layer |  | Outer layer
Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central | Side | Central
0.2 | 0.2 | 0.8 | 0.8 | 0.2 | 0.2 | 0.8 | 0.8 | 0.2 | 0.2 | 0.8 | 0.8 | 0.2 | 0.2 | 0.8 | 0.8