Magnets and Superconductors Group	August 2003

AT-MAS	Internal Note 2003-XX

### Field modulus calibration of the rotating coil measuring system for magnetic measurements at room temperature

S. Pauletta, AT-MAS

_____________________________________________________

#### Summary

The magnetic measurements at room temperature performed to monitor the LHC dipole production are based on the rotating coil technique. In this work, we describe a procedure to calibrate the measurements of the main field modulus using the Nuclear Magnetic Resonance (NMR). Results relative to the calibration campaign carried out in July-August 2003 of all measuring systems presently used at the manufacturer premises are presented.

Distribution:

AT-MAS Scientific Staff + AT-MTM Scientific Staff, P. Lebrun, N. Siegel

##### 1 Magnetic measurement systems

To detect assembly errors or faulty components at an early stage of production and to assess the field quality featured by collared coils and cold masses, magnetic measurements are performed at room temperature at the manufacturer premises. For this purpose, a provisional measuring system (called QUIMM) has been used for the pre-series, and a final one (called DIMM) is being delivered to the manufacturers. Both measuring systems adopt the rotating coil method.

- The QUIMM system uses one rotating coil mole to measure both assembly apertures. The measurement is performed along the axis at fixed positions, determined through fixed length of the shafts inserted along the bore tube. The mole can then rotate by means of an external engine, connected to it through the shafts.

- The DIMM system consists in two rotating coil moles, which are used to measure both apertures at the same time at fixed positions, where moles are placed by an automatic system. Moles then rotate by means of dedicated motorizations stacked to each mole and travelling with them.

##### 2. Calibration procedures

A magnetic field is measured both with the system to be calibrated and with a system of reference based on the Nuclear Magnetic Resonance method (NMR). In order to provide a magnetic field measurable by NMR, several 0.05 T (at 3 Amps) dipoles have been built. They all have the same design: an 88 cm long electro-magnet, whose vertical iron poles provide a uniform horizontal dipole field in a 6 cm aperture. A circular bore tube has been placed in the reference magnet to have the same lay out of the LHC main dipole aperture.

The magnet is provided with an NMR probe that measures the field intensity placed on its longitudinal centre. Since this probe measures the field intensity BNMRfat the central position only, while the harmonic coil system measures the field integral Bmisalong the central 750 mm of the dipole, the reference magnet must be mapped before the comparison between the two measurements can be performed. Therefore, a mobile NMR probe is used to measure BNMRm,i each 20 mm in i=1, 37 positions along the magnet axis. The rate:

ri=BNMRm,iBNMRf

is then computed for each of the 39 mapping positions between the mobile NMR probe and the fixed one (the first and the last measurement positions are only 1 cm spaced in order to evaluate extremity effects). The average <r> of all ri is then evaluated. This provides the ratio between the integrated field and the fixed NMR probe measurement.

Once the reference dipole has been mapped, it is then measured by the harmonic coil system to be calibrated. The field intensity featured by the reference magnet is higher than that featured by cold masses or collared coils during warm measurements. The mole rotation speed, then, must be decreased during the calibration, according to the sensitivity featured by the layout of the system electronics (amplifier gain, integrators, etc…), which is different between QUIMM and DIMM systems. The integrated value Bmismeasured by the system to be calibrated is then compared to the fixed NMR probe valueBNMRfmultiplied by the factor <r> to take into account of the difference between local and integral field:

Btrue=BNMRf<r>.

The difference between Bmis and Btrue can be caused by several reasons, such as the electronics of the measuring system, or the mole properties. In this second case the important parameter is the mole coil area, which is used to compute the magnetic field intensity from the magnetic flux. The new calibration results could be implemented in the system by changing the coil area (CAnew) with respect to the used one (CAold) by the following equation:

CAnew=BmisBtrueCAold.

##### 3. Calibration accuracy

The required accuracy for the feedback on the production is of 3.10-4 times the main field, i.e. 3 units. This accuracy is determined by the tolerances on the random component of the main field imposed by beam dynamics, which are of 8 units at one sigma. If the spread of the main field is larger than 8 units corrective actions can be taken using additional ferromagnetic laminations.

The accuracy featured by the calibration method described in the present work is mainly affected by the reference magnet performance, since the NMR measurement accuracy is of the order of 10-7 T. Even if DC powered with an high precision power supply, the field featured by the reference magnet is not perfectly stable during its heating up. To take into account such instabilities, the reference dipole mapping (i.e., the determination of the factor <r> should precede and follow the coil surface measurement at constant current (usually performed at 3 A) and the mean of the two mappings is then taken. The averaged rate <r> can vary, in fact, of about 10 units with respect to the first measured value, if both BfNMR and BNMRm,i rise in the first hours after the magnet has been switched on. To perform the calibration, then, the operator should wait the complete warming-up of the magnet or at least a time long enough to stabilize the magnet performance during the time needed for the calibration. In such a way, the magnet rate <r> can be measured with an accuracy of ±3 units (at 3 rms). The DC current used to power the reference magnet must be known with an accuracy of at least 10 units.

##### 4. Calibration results

The data coming from the mapping of the reference magnets used at the manufacturer premises are given in appendix A. The calibration results are given in Table 1 for each mole. The main result is that all systems, with the exception of the DIMM in Alstom, have a good calibration within 6 units. The DIMM system in Alstom shows a difference of -30 units; this effect is due to the electronics and up to now it has been corrected by applying -38 units of correction based on the results of the collared coil 1026 that has been measured with both DIMM and QIMM system. According to our calibration, the difference between DIMM and QIMM systems in Alstom is -34 units. Therefore, also in this case the calibration agrees with previous measurements within 6 units. When calibrating in Noell (both Zeisz and Wurzburg) and Ansaldo, an additional spare mole (number 18) has been brought and calibrated to check the reproducibility of the procedure. The result is very good, since the difference between the three measurements (and the calibration previously done at CERN) has been always within 1 unit.

##### 5. Conclusions

The systems used at the manufacturer premises to measure magnetic field at room temperature have been calibrated in July-August 2003. The procedure has been described, and the results show that discrepancies up to a maximum of 6 units are found. These results will be used together with information from measurements at 1.9 K to decide if a corrective action to reduce the spread in main field between firms is required.

##### Acknolwedgements

I would like to thank J. Billan, P. Galbraith, D. Giloteaux, and V. Remondino for having me introduced to the rotating coil measuring technique and for the constant support in these measurements. I also would like to thank E. Todesco and W. Scandale for supporting this work and useful discussions. A special thank to the personnel of Alstom, Ansaldo and Noell for their kind help.

##### Appendix

For each reference magnet used to calibrate the systems for the magnetic measurements at room temperature, the following figures show the profile of the factor ri:

ri=BNMRm,iBNMRf,

whereBNMRm,i and BfNMR are the field modulus measured by the mobile NMR probe and by the fixed one at each mapping position i. The reference field BfNMR is also reported for each mapping, together with the mean rate <r> used in the computations. From such plot, we can argue that the mean ratio <r> depends on the central field intensity, and becomes <r> higher when the central field is higher.

In all cases, more than one mapping has been performed in order to assess the magnet performance stability. In some cases, the average of the mean rate <r> has been considered into the computation, in some others the rate <r> has changed so much that some mappings has been discarded or used for some moles only.

| Dipole Mole # | 4 | 5 | 14 | 10 | 13 | 6 | 9 | 15 | 11 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| System | DIMM | DIMM | QUIMM | DIMM | DIMM | QUIMM | QUIMM | QUIMM | QUIMM |
| Place | Noell (Zeitz) | Noell (Zeitz) | Alstom | Alstom | Alstom | Ansaldo | Ansaldo | Noell (Wurzburg) | Noell (Wurzburg) |
| Date | 29-07-03 | 29-07-03 | 23-07-03 | 23-07-03 | 23-07-03 | 06-08-03 | 06-08-03 | 20-08-03 | 20-08-03 |
| Ref. Magnet | 900119 | 900119 | 900118 | 900118 | 900118 | 900117 | 900117 | 900121 | 900121 |
| <r> | 0.999081 | 0.999081 | 0.997586 | 0.998716 | 0.998716 | 0.999532 | 0.999532 | 0.997718 | 0.997718 |
| BNMR | 0.051842 | 0.051989 | 0.051771 | 0.051974 | 0.051966 | 0.051695 | 0.051696 | 0.051765 | 0.051765 |
| Bmis | 0.051781 | 0.051952 | 0.051663 | 0.051753 | 0.051740 | 0.051694 | 0.051703 | 0.051658 | 0.051669 |
|  | 0.051781 | 0.051953 | 0.051663 | 0.051752 | 0.051740 | 0.051695 | 0.051698 | 0.051658 | 0.051669 |
|  | 0.051781 | 0.051952 | 0.051664 | 0.051752 | 0.051740 | 0.051693 | 0.051699 | 0.051658 | 0.051668 |
|  | 0.051781 | 0.051953 |  | 0.051752 | 0.051740 | 0.051692 | 0.051707 | 0.051658 | 0.051669 |
|  | 0.051781 | 0.051952 |  | 0.051752 | 0.051740 | 0.051694 | 0.051702 | 0.051658 | 0.051669 |
| <Bmis> | 0.051781 | 0.051952 | 0.051663 | 0.051752 | 0.051740 | 0.051694 | 0.051702 | 0.051658 | 0.051669 |
| Btrue | 0.051794 | 0.051941 | 0.051646 | 0.051907 | 0.051899 | 0.051671 | 0.051672 | 0.051647 | 0.051650 |
| Diff. (units) | -2.5 | 2.2 | 3.4 | -29.9 | -30.7 | 4.4 | 5.8 | 2.2 | 3.7 |
| CAold | 3.72961 | 3.69538 | 3.68987 | 3.74265 | 3.72666 | 3.69858 | 3.72938 | 3.70457 | 3.74428 |
| CAnew | 3.72865 | 3.69618 | 3.69111 | 3.73147 | 3.71522 | 3.70021 | 3.73154 | 3.70537 | 3.74565 |

Table 1. Calibration results reported for each harmonic coil mole at the manufacturing premises. The difference between Bmis and Btrue is reported in units of 10-4.

Figure 1: Local values of the rate BNMRmBNMRf as coming from the mapping of the reference magnet 900118, placed in Alstom. The fixed NMR value and the average coefficient are also reported.

Figure 2: Local values of the rate BNMRmBNMRf as coming from the mapping of the reference magnet 900117, placed in Ansaldo. The fixed NMR value and the average coefficient are also reported.

Figure 3: Local values of the rate BNMRmBNMRf as coming from the mapping of the reference magnet 900119, placed in Noell (Zeitz). The fixed NMR value and the average coefficient are also reported.

Figure 4: Local values of the rate BNMRmBNMRf as coming from the mapping of the reference magnet 900121, placed in Noell (Wurzburg). The fixed NMR value and the average coefficient are also reported.