Description of the problem:  Referring to the email send by F. Savary today, the DIMM 4 measurement system for the collared coils broke down this morning. Due to the situation described, Ansaldo might soon face a shortage of collared coils which could result in a total stop of the production of this CMA. Therefore, we kindly ask for an urgent repair within the next two days.

Please consider also that:

- The last inspection of the measurement system has been carried out at the beginning of October 03. Since then, no further maintenance has been carried out. Maintenance each 3 months is considered necessary to guarantee a proper functioning of the measurement system as was agreed end of 2002.

- A maintenance cycle should at least include:

  - Exchange the moles with newly calibrated ones; in this case the pick-up surfaces given in the measurement software have to be adopted and the mupi-card has to be set or exchanged.

  - Check the state of the stainless steel cable and the rolls.

  - Check the twist of the connection cable, its integrity and the fixation of the extensions.

  - Eventually the stored files in the sun station have to be compressed in order to keep storage space for the next measurements.

  - Check and eventually update the software version, if necessary.

- Ansaldo also repeatedly reported problems with the stiffness of their cables which causes problems for the mole leveling.

- Exchange of the feet of the mole to Delran feet.

- Please check the brass sliding rings on the motor mole if they are very used and if the screws that are fixing the rings start to protrude (as have been observed in other companies). If so, the brass rings should be exchanged as well, in order to avoid scratches from the protruding screws inside the cold bore tube.

# Text boxes

Request for Maintenance/Repair

Date: 22-March-04

For: DIMM System 4; used for warm magnetic measurements on collared coils

At: Ansaldo, Genova

To: AT-MTM-FM, Juan Garcia Perez, David Giloteaux, Dominique Cote, Adriaan Rijllart

From: Christine Vollinger, AT-MAS

CC: M. Bajko, A. Musso, F. Savary, L. Rossi, W. Scandale, L. Walckiers, J. Vlogaert.