Description of the problem: During the visit of the Noell premises on 17th -18th Feb 2004 in Wuerzburg, Noell personnel indicated the following problems with their measurement system:

- The last inspection of the measurement system has been carried out by P. Galbraith the 07th Oct 03. Since then, no further calibration or maintenance has been carried out. Maintenance each 3 months seems necessary to guarantee a proper functioning of the measurement system as was agreed end of 2002.

- A maintenance cycle should at least include:

  - Exchange the moles with newly calibrated ones; in this case the pick-up surfaces given in the measurement software have to be adopted.

  - Check the state of the stainless steel cable and the rolls.

  - Check the twist of the connection cable, its integrity and the fixation of the extensions.

  - Eventually the stored files in the sun station have to be compressed in order to keep storage space for the next measurements.

  - Check and eventually update the software version, if necessary.

- The disk drive of the sun station in the measurement rack is blocked and should be exchanged. The disk drive is needed to carry out the software updates on the system and to load the tables in the case of special measurements.

- Exchange of the feet of the mole to Delran feet.

- The brass sliding rings on the motor mole are very used and the screws that are fixing the rings started to protrude. In order to avoid scratches from the protruding screws inside the cold bore tube, Noell personnel has already manually filed down the heads of these screws. The brass rings should be exchanged fast (see picture below).

- The cables show signs of ageing. Eventually the nylon protection has to be replaced.

# Text boxes

Request for Maintenance/Repair

For: DIMM System 3; used for warm magnetic measurements on collared coils

At: Noell, Wuerzburg

To: AT-MTM-FM, Juan Garcia Perez, David Giloteaux, Dominique Cote

From: Christine Vollinger, AT-MAS

CC: P. Fessia, J. Miles, V. Remondino, L. Rossi, W. Scandale, L. Walckiers, J. Vlogaert.