MEMORANDUM

31th August 2007

EDMS n. xxxxxx

Subject: Correction of the MQY cold mass field quality measurements

From: P. Hagen, E. Todesco

To: L. Bottura, L. Deniau, J. P. Koutchouk, N. Catalan-Lasheras, N. Sammut, S. Sanfilippo

The issues

During the production of the MQY quadrupoles for the insertion regions it was discovered an unusual spread in the transfer function. The next two magnets were built and tested to check the effect at “warm” and “cold”, with no effort at sorting the collars by permeability level. The effect at “warm” was confirmed to disappear at “cold”. It was further found that collar permeability was significantly higher in these magnets. It was decided to measure the collar permeability for all remaining apertures to be built. Unfortunately there are no permeability measurements from the start of the MQY production [1].

Table I gives a summary of the MQY warm measurements after the production has ended. It shows 3 interesting features:

- There are permeability measurements for 14 out of 60 apertures. The value given represents an average for the magnet as there is some variability in the same collaring pack.

- The production information for MQY 19 in MTF [2] contain no cold mass measurements

- The production information reports that aperture that aperture 41 is present in both magnet 21 and 29

Table II gives a summary of the MQY cold measurements after the testing at CERN has ended. It shows 2 interesting features:

- 6 out of 18 cryostats have no cold field quality measurement

- 6 cryostats consist of two MQY magnets in series

If all the MQY magnets had been tested at cold conditions the warm data would not have been needed. However, it seems there are 3 good reasons why the warm data is of interest beyond the trivial case of using it for cross-checking:

- 6 cryostats need to be extrapolated at cold using a warm-to-cold correlation

- The cold measurements can only gives integrals for the entire cryostat. Our AB/ABP colleagues would prefer to have the field quality information per magnet, which is the unit of resolution in the MAD-X description of the LHC accelerator [3].

- There is some calibration issue with the cold measurement system [4] concerning the transfer function, and the warm data is used for finding the correct offset.

The remaining of this memo deals with the correction of the measurements affected by the elevated permeability.

Table I. Summary of MQY “warm” measurements

Table II. Summary of MQY “cold” measurements

The warm measurement data

Plots of the original warm data are shown in Fig. 1 to 4.

The normalised G/i in Fig. 1 is based upon the assumption that the measurement current was 1.5 A for aperture 1 and 3.0 A for the others. The current is not recorded in the data. Based upon experience from similar measurements with the MQ type, the deviation in desired and actual excitation current could be in the order of 1 mA giving an uncertainty of ~ 5 units. The G/i are constant for apertures 9 to 17. When we compare these values to the integrated G/i we see a variation as expected and assumes that the measurement system records the G/I dl integral correctly, but the natural variation in G/i has wrongly been attributed to the magnetic length for apertures 9 to 17. Since these apertures have no associated permeability measurement, they will not complicate the treatment in this memo.

Fig. 1. Measured G/i in the cold mass

Fig. 2. Measured integrated G/i dl in the cold mass

Fig. 3. Measured b6 in the cold mass

Fig. 4. Measured b10 in the cold mass

The statistics (mean and sigma) of the data in Fig. 1 to 4 is shown in Table III. The spread of the Gdl/i being smaller than the spread of the G/i confirms the previous working assumption that Gdl/I was the quantity actually measured, but not always correctly split between magnetic length and normalised gradient. Judging the spread of the integrated transfer function against other similar magnet families like MQ and MQM we would expect a value between 10 and 15 units. The 23 units observed is indeed a significant quantity.

Table III. Summary of MQY “warm” statistics

The permeability sensitivity

We assume that the elevated permeability changes the transfer function and allowed harmonics as linear functions as shown in Eq. 1 and 2 where S are signed sensitivity coefficients.

G/i’cm = G/icm + Sg ·μr									           (1)

bn’cm = bncm + Sn· μr									 	(2)

The S coefficients should normally have been computed from the magnetic model input to BEM-FEM code [5]. The S coefficients that were available did not match the observed effect. Therefore we pursued the idea of guessing the missing permeability values based upon the measurement data. This is based upon the following assumptions about its effect:

- leaves a strong signal compared to all other sources of variability

- perturbs simultaneously the gradient and the harmonics b6 and b10

Plots of the original warm data versus permeability are shown in Fig. 5 to 7. The straight line represents a linear fit of the data and from the equations given we end up with a syntetic sensitivity table shown in Table IV. Evidently we see that not all variability can be due to the permeability but this should be largely cancelled out by the linear fit function.

Fig. 5. Measured G/i in the cold mass vs measured relative permeability

Fig. 6. Measured b6 in the cold mass vs measured relative permeability

Fig. 7. Measured b10 in the cold mass vs measured relative permeability

Table IV. Sensitivity table constructed from linear fits

The S coefficients in table IV are to be understood as follows. If the relative permeability change with Δμr = 1 (an “astronomic” change), then the normalised gradient will increase by almost 50% whilst b6 will decrease with -1.4% and b10 hardly change.

Permeability estimates

We know take the subset of the measurement data, namely <μr  G/i  b6 b10 > and split them into one table with known permeability (because measured) and one table with unknown permeability.

Furthermore we use the inverse functions f() → Δμr   for each quantity in table with known permeability. Each of these 3 quantities is weighted with the inverse of its relative sigma to produce a final guess. The remaining difficulty is to add a constant offset we do not know precisely, namely the nominal μr representing the normal permeability for the unperturbed collars. We estimate this value by using a quadratic error function formed by the sum of the ε between permeability measured and guessed. This function has a global min around 1.003 as shown in Fig 8. This is a value that confirms with typical values measured for the MQ magnet family.

The computation of the expected μr for the table with unknown permeability uses the same inverse functions, weights and nominal estimate for permeability. After some tedious treatment we end up with a μr value for each aperture.

Correction of the cold mass measurements

The last part of the exercise is to correct the cold mass measurements for the perturbed permeability. We know that the effect should almost or mainly disappear at cold so we just use the sensitivity table to adjust the warm data by removing the estimated effect. The treatment is straightforward and the results are shown in Fig. 9 to 11 using the same scale on the vertical axis as in Fig. 2 to 4. The statistics summary is shown in Table V and should be compared with Table III. The spread in the integrated transfer function has dropped from 23 units to 12 after the correction. This spread is similar to the ones known for the MQM and MQ magnet families. This is a cross-check that the correction has been “reasonable”.

Fig. 8. “Global” relative mu error as function of nominal relative mu

Fig. 9. Measured integrated G/i dl in the cold mass (corrected for elevated μ)

Table V. Summary of MQY “warm” statistics (corrected for elevated μ)

Fig. 10. Measured b6 in the cold mass (corrected for elevated μ)

Fig. 11. Measured b10 in the cold mass (corrected for elevated μ)

The “warm to cold” correlation

The most interesting feature of the “warm” data is how they can be used to predict the field quality at “cold” operational conditions. The following figures (Fig. 12 to 19) show the plots of warm-cold correlation before and after the data has been corrected for elevated μ. The plots have not changed significantly but the spread in the horizontal axis (warm) is reduced. This is most clearly visible in Fig. 19.

Fig. 12. G/i  correlation for “warm to cold”

Fig. 13. G/i  correlation for “warm to cold” (corrected for elevated μ)

Fig. 14. G/I dl  correlation for “warm to cold”

Fig. 15. G/i dl  correlation for “warm to cold” (corrected for elevated μ)

Fig. 16. b6  correlation for “warm to cold”

Fig. 17. b6  correlation for “warm to cold” (corrected for elevated μ)

Fig. 18. b10  correlation for “warm to cold”

Fig. 19. b10  correlation for “warm to cold” (corrected for elevated μ)

Conclusion

Table VI shows the cold-warm offsets before and after correction. The μ-correction seems to be effective although the b6 at injection did not improve much, but the improvement in the G/I dl is significant and similar to the one found for the MQM and MQY magnet families.

Table VI. Average and standard deviations for cold-warm offsets, before and after μ-correction

Acknowledgements

We wish to thank our colleagues in AT and ABP for pointing out the anomalies in the magnetic measurements and subsequent discussions.

References

[1] 	R. Ostojic, private communication

[2] 	MTF. Manufacturing and Test Folders Travellers, 
web url https://edms.cern.ch/asbuilt/plsql/mtf.home

[3]	M. Giovanozzi and F. Schmidt, private communication

[4]	E. Todesco, private communication

[5] 	S. Russenschuck, CERN Yellow Report 99-01, ROXIE: Routine for Optimization of Magnet 
X-Sections, Inverse Field Calculation and Coil End Design